Each sealed source contained in a portable gauging device must be tested at regular intervals to ensure that the radioactive material is secure within its capsule and not leaking contamination.  Leak test (LT) requirements are specified in section 64E-5.1303, Florida Administrative Code.

I.	Leak Test Frequency

- CPN International, Inc. gauges will be leak tested at least every 12 months.

- Troxler Electronic Laboratories, Inc., Model 3216, 3241-C, 3241-G, 3242, 3430, 3430-M, 3440, 3440-M, 3450, 3451, 4232, 4301, 4302, 4430, 4640, 4640-B gauges will be leak tested at least every 12 months.

- All other portable gauges will be leak tested at least every 6 months.

II.	Leak Test Kit

Only LT kits provided by licensed vendors will be used to sample (smear) sealed sources contained in portable gauging devices.

III.	Taking the Leak Test Sample

LT samples will be taken only by Authorized Users, wearing their assigned personnel monitoring badges.  LT samples will be taken in accordance with the written instructions provided by the supplier of the LT kit and the portable gauging device manufacturer.

IV.	Leak Test Sample Analysis

LT sample analysis will be performed only by vendors specifically licensed to provide the service by the Florida Bureau of Radiation Control, the U.S. Nuclear Regulatory Commission, or other state radiation control agencies.

V.	Leak Test Records

If a test indicates a device’s sealed source is contaminated, the portable gauging device will be removed from service and the Florida Bureau of Radiation Control will be notified immediately at (407) 297-2095.  A written report on the leaking source will be submitted to the bureau within 5 days (mailing address: Florida Bureau of Radiation Control, Radioactive Materials Program, Bin C21, 4052 Bald Cypress Way, Tallahassee, FL 32399-1741).  The report will describe the equipment involved, the test results, and the corrective actions taken (i.e., device removed from service until repaired; radiation surveys conducted to determine presence of contamination; decontamination as necessary).

Leak test records will be retained for 3 years.  The records will include the following information:

- Each source’s manufacturer name, model, and serial number;

- The identity of each sealed source radionuclide and its estimated activity, expressed in microcuries (or becquerels);

- The measured activity of each leak test sample, in microcuries (or Bq);

- The date the sample was collected; and

- The signature of the Radiation Safety Officer (or the RSO’s designee).