Information Notice Number 12 (December 1997) For Florida Radiation Machine Registrants Department of Health, Bureau of Radiation Control

RADIATION PROTECTION PROGRAM FOR MAMMOGRAPHY REGISTRANTS

The Florida Administrative Code requires radiation machine registrants to develop, document and implement a radiation protection program (RPP).  A properly developed and implemented RPP can minimize unnecessary radiation exposure to employees and the general public, help prevent either group from exceeding their legal limits on exposure to radiation, and document such an effort has been made by the registrant.  The RPP requirement for mammography only facilities can be satisfied in one of the following three ways:

I. Pre-approved Radiation Protection Program: Mammography (see attachment).

The requirements for an RPP will have been met if:

a) If the conditions at the registrant’s facility satisfy line items 1 through 9,

b) the name of the person responsible for the radiation protection program is entered in line item 10,

c) the dates the plan is in effect [Note: a plan is valid at most for one calendar year, it must be reviewed and dated in January of every year, see line item 7] i.e., the current month through December of the current year, are entered in line item 11,

d) the facility name and x-ray machine registration number are entered at the bottom, and

e) the form is signed and dated by the responsible party.

This plan does NOT have to be submitted for approval.

II. Alternate Radiation Protection Program (see attachment).

The alternate plan offers the registrant a structured set of options for developing an RPP.  For each numbered step one and only one option must be selected.  If the second choice in a step is selected, the required documentation must be attached to the plan.  This plan and the required attachments MUST be submitted to the program office for review.  Previously reviewed alternate radiation protection program do not need to be resubmitted for review unless changes are made.

III. Self-developed Radiation Protection Program.

The registrant may develop an RPP to fit their specific situation.  The provisions of this RPP should address all requirements in the Florida Administrative Code.  Self-developed programs MUST be submitted to the program office for review.  Previously reviewed self-developed radiation protection program do not need to be resubmitted for review unless changes are made.

If you have questions or need guidance, please contact this office at:

Department of Health Bureau of Radiation Control, Radiation Machine Section 4052 Bald Cypress Way, Bin C21 Tallahassee, FL 32399-1741 Phone: (850) 245-4888 Fax: (850) 617-6442 www.myFloridaEH.com/radiation

PRE-APPROVED RADIATION PROTECTION PROGRAM

MAMMOGRAPHY

Note: Radiation Protection Programs are used to minimize unnecessary exposure to employees and the general public and help prevent either group from exceeding their legal limits on exposure to radiation.  This program has been developed by the Bureau of Radiation Control to assist registrants in meeting Florida Administrative Code requirements.  If this program is used, it does not need to be submitted for review.

The signature below attests this registrant has evaluated the following provisions and agrees they accurately describe the conditions present.  All staff members involved in making radiographic exposures will be made aware of these provisions and held accountable for them.

1.	X-ray machine operators will remain completely behind a fixed barrier during all radiographic exposures.

2.	Only patients are allowed in the x-ray room and patients are not routinely held during radiographic exposures.  If under unusual circumstances a patient must be held:

a)	A log will be kept documenting the date, the name of the individual holding the patient, type of procedure and number of films taken.

b)	The holder will be positioned so no portion of their body is in the primary x-ray beam.  The only exception may be hands when necessary. In those cases, lead gloves will be required.

c)	The holder will be protected from scattered radiation with a lead apron.

d)	If patients are held more three times in any calendar quarter, holding of patients will be discontinued until an alternate plan is developed by the registrant and approved by the Bureau of Radiation Control, Radiation Machine program.

3.	Personnel monitoring records are reviewed on receipt to ensure no staff member exceeds 100 mR in the calendar year.  If any staff member exceeds 100 mR an Alternate Plan will be developed and submitted to the Bureau of Radiation Control, Radiation Machine Program for approval.

4.	Fluoroscopic exposures are never made.

5.	Declared Pregnant Women never take radiographic exposures or when a woman declares pregnancy in writing (including estimated conception date) they are provided with a monitoring device specifically for documenting fetal dose.  The device will be worn at waist level and underneath any protective clothing.

6.	The radiation protection program will be evaluated annually during the month of January for accuracy.  An attached sheet will be used to document this review.  Each review will include a date of review and the signature of the reviewer.

7.	The radiation protection program will be re-evaluated when changes occur that could affect dose.

8.	All radiation protection program(s) will be kept on file until the registration is terminated.

9.	________________________________________ is responsible for radiation protection.

10.	The above provisions are true from ____________________ through December of ________.
						(Month/Year)					(Year)

|  |  | JR |
| --- | --- | --- |
| Facility Name (print or type) |  | X-ray Machine Registration Number |
| Signature of Person Responsible for Above Program |  | Date Signed |

ALTERNATE RADIATION PROTECTION PROGRAM

Note: Radiation Protection Programs (RPPs) are used to minimize unnecessary exposure to radiation workers and the general public and help prevent either group from exceeding their legal limits on exposure to radiation.  This program has been developed by the Bureau of Radiation Control to assist registrants in developing an RPP which meets the Florida Administrative Code requirements.  This program and all required attachments must be submitted to the Bureau of Radiation Control, Radiation Machine Program for review.

The signature below attests this registrant has evaluated the following provisions.  One option for each numbered section has been selected and accurately describes the conditions present at this facility.  All staff members involved in making radiographic exposures will be made aware of these provisions and held accountable for following them.

RADIOGRAPHIC EXPOSURES

1.	_____	X-ray machine operators will remain completely behind a fixed barrier during all radiographic exposures.

_____	Written procedures are followed describing how aprons and movable barriers are used when operators are not positioned behind a fixed barrier during radiographic exposures.  These procedures are attached and are part of this Alternate Plan.  In addition, surveys, calculations or monitoring records are attached to document members of the public and occupationally exposed persons do not exceed dose limits when the operator is not completely behind a fixed barrier during radiographic exposures.

2.	_____	Patients are not held during radiographic exposures.

_____	Written procedures are followed describing how aprons and movable barriers are used when patients are held during radiographic exposures.  These procedures are attached and are part of this Alternate Plan.  In addition, holding logs or monitoring records are kept to document members of the public and occupationally exposed persons do not exceed dose limits when any person other than the patient is allowed in the x-ray room during radiographic exposures.

3.	_____	Only patients are allowed in the x-ray room during radiographic exposures.

_____	Written procedures are followed to assure that any person in the x-ray room during radiographic exposures is medically necessary.  Written procedures are followed that describe how aprons and movable barriers are used when any person other than the patient is allowed in the x-ray room during radiographic exposures.  These procedures are attached and are part of this Alternate Plan.  In addition, surveys, calculations or monitoring records are attached to document members of the public and occupationally exposed persons do not exceed dose limits when patients are held during radiographic exposures.

4.	_____	Less than 100 radiographic exposures are made per week when averaged over a year or the portion of the year the facility is open for business.

_____	Surveys, calculations or monitoring records are attached to document members of the public do not exceed dose limits outside any radiographic x-ray room where 100 or more radiographic exposures are made per machine each week.

RADIOGRAPHIC EXPOSURES (continued)

5.	_____	Any wall bucky or cassette stand is located on an outside wall or a wall with at least one sixteenth (1/16) inch of lead equivalent shielding.

_____	Surveys, calculations or monitoring records are attached to document members of the public do not exceed dose limits outside any radiographic x-ray room where the wall bucky or cassette stand is not located on an outside wall nor a wall with at least one sixteenth (1/16) inch lead equivalent shielding.

6.	_____	The center of any radiation machine table is at least two (2) feet from any wall.

_____	Any wall less than two (2) feet from the center of the radiation machine table is either an outside wall or a wall with at least one thirty-second (1/32) inch lead equivalent shielding.

_____	Surveys, calculations or monitoring records are attached to document members of the public do not exceed dose limits outside any radiographic x-ray room where the center of the radiation machine table is not at least two (2) feet from any inside wall and the wall does not have at least one thirty-second (1/32) inch lead equivalent shielding.

FLUOROSCOPIC EXPOSURES (if applicable)

7. 	_____	Fluoroscopic machine operators remain completely behind a fixed barrier during all fluoroscopic exposures.

_____	Written procedures are followed describing how aprons and movable barriers are used when operators are not positioned behind a fixed barrier during fluoroscopic exposures.  These procedures are attached and are part of the Alternate Plan.  In addition, surveys, calculations or monitoring records are attached to document members of the public and occupationally exposed persons do not exceed dose limits when the operator is not positioned behind a fixed barrier during fluoroscopic exposures.

8. 	_____	Patients are not held during fluoroscopic exposures.

_____	Written procedures are followed that describe how aprons and movable barriers are used when patients are held during fluoroscopic exposures.  These procedures are attached and are part of this Alternate Plan.  In addition, holding logs or monitoring records are kept to document members of the public and occupationally exposed persons do not exceed dose limits when patients are held during fluoroscopic exposures.

9. 	_____	Only patients are allowed in the x-ray room during fluoroscopic exposures.

_____	Written procedures are followed to assure that any person in the x-ray room during fluoroscopic exposures is medically necessary.  Written procedures are followed that describe how aprons and movable barriers are used when any person other than the patient is allowed in the x-ray room during fluoroscopic exposures.  These procedures are attached and are part of this Alternate Plan.  In addition, surveys, calculations or monitoring records are attached to document members of the public and occupationally exposed persons do not exceed dose limits when any person other than the patient is allowed in the x-ray room during radiographic exposures.

FLUOROSCOPIC EXPOSURES (continued)

10. 	_____	Fluoroscopic x-ray machines are operated less than fifteen (15) minutes per machine each week.

_____	Surveys, calculations or monitoring records are attached to document members of the public dose limits are not exceeded outside each fluoroscopic x-ray room where fluoroscopic x-ray machines are operated fifteen (15) or more minutes per machine each week.

11. 	_____	High level (>10 R/min) fluoroscopy is not used at this facility.

_____	Surveys, calculations or monitoring records are attached to document members of the public dose limits are not exceeded outside each room where high level fluoroscopy is used.

GENERAL REQUIREMENTS

12.	_____	Minors never make radiographic and fluoroscopic exposures.

_____	Written procedures are followed when minors make radiographic or fluoroscopic exposures.  These procedures are attached and are part of this Alternate Plan.

13.	_____	Declared Pregnant Women never make radiographic and fluoroscopic exposures.

_____	When a woman declares her pregnancy or intent to become pregnant in writing (this statement should include estimated conception date) they are provided with a monitoring device specifically for documenting fetal dose.  The device will be worn at waist level and underneath any protective clothing.

14.	The provisions of this radiation protection program will be evaluated annually during the month of January to ensure they accurately describe the conditions expected in the coming year.  An attached sheet will be used to document this review.  Each review will include a date of review, the signature of the reviewer, and the statement the provisions are true for the coming January through December.  If the program needs to be changed, the revisions will be sent to the Bureau of Radiation Control, Radiation Machine Program for review.

15.	This radiation protection program will be re-evaluated when changes occur that could affect dose to radiation workers or the general public.

16.	All radiation protection program(s) will be kept on file until the registration is terminated.

17.	________________________________________ is responsible for radiation protection.

18.	The above provisions are true from ____________________ through December of ________.
					  (Month/Year)				       (Year)

|  |  | JR |
| --- | --- | --- |
| Facility Name (print or type) |  | X-ray Machine Registration Number |
| Signature of Person Responsible for Above Program |  | Date Signed |