Regulatory Guide 1.20
Revision 2
Issuance Date: February 2008

Instructions for Preparing Applications
for Radioactive Materials Licenses Authorizing

## USE OF SEALED SOURCES
IN FIXED GAUGING DEVICES

TABLE OF CONTENTS

Page

I.	INTRODUCTION	1-4

A.	Purpose of Guide	3-4

B.	Applicable Regulations	3

C.	License Requirements and Restrictions	4

II.	FILING AN APPLICATION	5-6

A.	General	5

B.	License Fees	6

III.	CONTENTS OF AN APPLICATION	6-16

Item 1.a.	Name and Mailing Address of Applicant	6

Item 1.b.	Street Address(es)	7

Item 2.a.	License Fee Category	7

Item 2.b.	License Fee Enclosed	7

Item 3.	Application Type	7

Item 4.	Individual Users	7

Item 5.	Radiation Safety Officer	7

Item 6.	Training and Experience in Radiation Safety	8

Item 7.	Radioactive Material	8

Item 8.	Use of Radioactive Material	8

Item 9.	Radiation Detection Instruments	9

Item 10.	Instrument Calibration	10

Item 11.	Personnel Monitoring Devices	10

Item 12.	Facilities and Equipment	11

Item 13.	Radiation Protection Program	12-16

Item 14.	Waste Disposal	1

Item 15.	Certificate	15

Item 16.	*NEW - Increased Security Controls – If applicable       	16

IV.	LICENSE AMENDMENTS	17

V.	LICENSE RENEWAL	17

VI.	LICENSE TERMINATION	17

TABLE OF CONTENTS

APPENDICES

App. A.	Member of the Public (MOP) Dose Limit Compliance Study

App. B.	ALARA Policy

App. C.	Duties and Responsibilities of the Radiation Safety Officer

App. D.	Radiation Safety Training Program

App. E.	Operating Procedures

App. F.	Personnel Monitoring Procedures

App. G.	Procedure for Release From Personnel Monitoring Requirements

App. H.	Posting Procedure

App. I.	Lock-out Procedure

App. J.	Emergency Procedures

App. K.	Procedures for Ordering, Receiving, Opening and Shipping Packages Containing Radioactive Material

App. L.	Leak Testing Procedure

App. M.	Inventory Procedure

App. N.	Record Retention Procedure

App. O.	Notification and Reporting Procedure

App. P.	Transfer/Disposal Procedure

EXHIBITS

Ex. A	Inventory Form

Ex. B.	Survey Form

Ex. C.	Shipping Paper – Private Use Shipment

Ex. D.	Emergency Response Information

Ex. E.	Instructions for Women Working With Radiation

Ex. F.	Instructions to Declared Pregnant Women

Ex. G.	Declared Pregnancy Form

Ex. H.	Delegation of Authority to Make Legally Binding Statements

SUPPLEMENTS

Supp. A.	Form DH-1054 “Application for Radioactive Materials License – Non-Human Use”

Supp. B.	List of Attachments Provided With Application

Supp. C.	Fixed Gauging Device License Application Checklist

ii

I.	INTRODUCTION

A.	PURPOSE OF GUIDE

This guide provides instructions for preparing an application for a state of Florida radioactive materials license authorizing possession and use of sealed sources in fixed gauging devices – industrial gauging devices that use sealed radioactive sources, typically installed in fixed locations, as a process tool.  This guide also contains additional information that may be used as a training resource.  In addition, the model procedures are written to be stand-alone documents so several acronyms and abbreviation as well as other information may be repeated.  Fixed gauging devices are designed and manufactured for the purpose of determining or controlling a process parameter such as thickness, density, level, interface location, particle size distribution, or qualitative or quantitative chemical composition.  Nuclear gauging systems typically consist of:

- A sealed radioactive source – gamma, beta or neutron emitting isotope (e.g., cesium-137, cobalt-60, strontium-90, krypton-85, californium-252, etc.);

- A source holder (source head)* – the device used to support and retain the source, attenuate the primary radiation beam, and provide shielding to reduce scattered radiation;

- A radiation detector consisting of a sensor and amplifier that converts ionizing radiation into electrical impulses and amplifies the signal that is produced to provide parameter data; and

- An electronic component (usually computerized) that converts the amplified signal into a format that displays the data on a monitor or directly adjusts process parameters.

Process
 material

Primary beam

Source

Shutter

Shielding

Detector

Source holder
(source head)

##### Typical Fixed Gauging Device Design

Fill Level Gauge

Insertion Gauge

Flow Gauge

Thicknes Gauge

Fill Level Gauge

Density/Flow Gauge

Thickness Gauge

Insertion Gauge

##### Common Types of Fixed Gauges

Only the source and source holder are subject to licensure and regulatory control.

* The terms source holder and source head are not always used as synonyms; source holder is also used to refer to the component inside the source head that holds the source in place within the head.

License Types: General and Specific

Possession and use of fixed gauging devices in Florida requires a radioactive materials license issued by the state.  There are two license types: specific license and general license.

Gauging devices that present more significant health and safety hazards require a specific license.  The specific license is a document issued to an applicant, authorizing a particular use of a gauge.  The license describes the types of devices the licensee may possess and the limits imposed on their use.  Applicants must demonstrate that they have personnel qualified to work with the devices, and appropriate facilities, equipment and procedures to ensure safe operations.  The fees for a specific license are higher than for a general license.  The specific license application process is detailed, requiring establishment and implementation of a detailed radiation protection program.  The specific license application process is detailed, requiring establishment and implementation of a comprehensive radiation protection program.  Personnel operating under a specific license may be approved to perform “advanced” activities such as installations and relocations that are not allowed under a general license.

A general license does not require the filing of an application with the department or the issuance of a licensing document.  The general license becomes effective upon receipt of the radioactive material.  The distributor is required to notify Florida whenever radioactive material has been delivered to a Florida customer.  While general licenses offer ease of acquisition, minimal documentation requirements and lower fees, they are subject to certain conditions, requirements, limitations and restrictions.  For example, installation and relocation of some generally licensed fixed gauging devices can only be performed by trained individuals authorized under a specific license.  Some generally licensed fixed gauging devices may be mounted, but cannot be unlocked and put into use until a specifically licensed individual, such as a gauge manufacturer representative or an independent gauge services provider performs radiation measurements and tests the gauge’s alignment with the detector to ensure that it will operate safely.  Generally licensed gauges may also have restrictions on the types of non-routine maintenance that may be performed on them by the owner.  The label on a generally licensed gauge describes the restrictions that apply to that model.  For additional information concerning issuance of a general license or the various types of general licenses, review section 64E-5.204, F.A.C., and section 64E-5.206, F.A.C.

Florida allows authorization for a general license to be included under a specific license.  Applicants seeking to include authorization for generally licensed radioactive materials as part of their specific license must include a written request for general license authorization as part of their application; see section III.7. of this guide for additional guidance.  Once approved, authorization for generally licensed radioactive materials will be included as part of Items 6, 7, 8 and 9 of the license, and a general license condition describing requirements associated with generally licensed radioactive materials will be included in the license.  This condition does not provide an exemption from compliance with the current regulatory requirements.  Refer to Information Notice 2007-02, which summarizes additional requirements that were incorporated in Revision 6 to Part II, Subpart B, Chapter 64E-5, F.A.C., dated September 28, 2007.

This guide provides instructions for applicants seeking a fixed gauging device specific license.  Unless otherwise noted, when used in this guide, the term “license” refers to a specific license.

Appendices, Exhibits and Supplements

Applicants must acquire and maintain appropriate facilities and equipment, have appropriately trained workers, and implement procedures that ensure compliance with regulatory requirements.  This guide provides a set of appendices, exhibits and supplements to assist in the development of a radiation protection program.

- Appendices are model procedures that may be used to address regulatory requirements.

- Exhibits are samples of the types of documents or forms that must be submitted as part of the application, and in some cases, are model forms that may be used to satisfy regulatory requirements.

- Supplements include resources for preparing the application and additional resources and reference material.

Model procedures and forms may be adopted by submitting them as part of the license application, or may be used as guides for developing equivalent procedures and forms.  Carefully review the regulations, model procedures and forms before deciding if the models are appropriate for the activities being requested.

IMPORTANT NOTICE:

The information provided in a license application must demonstrate that proposed equipment, facilities, personnel and procedures are adequate to protect public health and property in accordance with regulatory requirements.  Submission of incomplete or inadequate information will result in delays in the license approval process.  Additional information will be requested when necessary to ensure that an adequate radiation safety program has been established.  Such requests will delay completion of the application review, and may be minimized by a thorough study of the regulations and this guide prior to submitting the application.

While the adoption of the attached model procedures and forms provides for a radiation protection program that complies with regulatory requirements, applicants should consider additional equipment, procedures and training that may be appropriate for the scope of their operations.

B.	APPLICABLE REGULATIONS

Florida is an Agreement State; it has an agreement with the U.S. Nuclear Regulatory Commission (NRC) to assume regulatory authority over most activities involving radioactive material within the state.  With certain exceptions, the Department of Health (department), Bureau of Radiation Control (bureau) regulates the possession and use of radioactive material within Florida.  Exceptions include nuclear power plants and federal agencies, and national security issues involving radioactive material, which remain under NRC jurisdiction.

Under authority of Chapter 404, Florida Statutes (the Florida Radiation Protection Act), the bureau issues licenses to users of radioactive material and performs inspections to ensure safe operations and compliance with Chapter 64E-5, Florida Administrative Code (F.A.C.), the department’s control of radiation hazard regulations.  Chapter 64E-5, F.A.C., is available on the Internet at http://www.doh.state.fl.us/environment/radiation.  The bureau amends these regulations periodically.  Licensees are notified of changes as they occur.  When applicable, licensees will need to revise their safety programs to address changes in regulatory requirements.

The parts of the regulations applicable to fixed gauging devices licensees are listed below, and should be used in conjunction with this guide.  Chapter 64E-5, F.A.C., Part XIII, Subparts A and C include rules that specifically address fixed gauges.

Part I	“General Provisions”

Part II	“Licensing of Radioactive Materials”

Part III	“Standards for Protection Against Radiation”

Part IX	“Notices, Instructions and Reports to Workers; Inspections”

Part XIII	“Radiation Safety Requirements for Possession and Use of Sealed or Unsealed 
	 Sources of Radioactive Materials”

Part XV	“Transportation of Radioactive Materials”

Fixed gauging device users are also subject to U.S. Department of Transportation (DOT) regulations, which are found in Title 49, Code of Federal Regulations (49 CFR), and are incorporated into Chapter 64E-5 by reference.  DOT regulations are available on the Internet at http://www.access.gpo.gov/nara/cfr/cfr-table-search.html#page1 and can be ordered from the U.S. Government Printing Office by calling (904) 353-0569 or writing 100 West Bay Street, Suite 100, Jacksonville, FL 32202, Attn: Superintendent of Documents.

C.	LICENSE REQUIREMENTS AND RESTRICTIONS

Licensees are required to confine use and possession of radioactive material to the locations and purposes authorized by the license.  The license is divided into two sections: Items and Conditions, which are described below.  The first section of the license lists Items 1 - 9.  The remainder of the document lists the license conditions, which may vary in number based on the authorizations provided by the license, but always begin with Condition 10.

License Items

| Item No. and Title | Description |
| --- | --- |
| 1. Name | Lists the legal name of the licensee (individual or business). If the license is issued to a business, Item 1 must list the company’s name as it is registered with the Florida Department of State, Division of Corporations; (850/488-9000 or http://www.sunbiz.org/. If a business operates under another name, Item 1 must list both the registered name and the fictitious name it is doing business as (d/b/a). |
| 2. Address | Lists the mailing address, which may be different from the physical address where records and gauges are used/stored. If the two addresses are different, the physical address must be listed in License Condition 10; if they are the same, Condition 10 references the address listed in License Item 2. |
| 3. License Number | Lists the number assigned to the license by the bureau. The number should be referenced in all correspondence with the bureau. |
| 4. Expiration Date | Lists the date the license will expire. A radioactive materials license is valid for 5 years from the date of issuance. Applications for renewal must be received 30 days prior to the expiration date for the license to be deemed “timely filed”. |
| 5. Category | Activities involving possession and use of radioactive materials are divided into license categories. Fixed gauging devices are covered under Category 3H. Conducting more than one category of licensed activity requires a separate license for each category of use. Section 64E-5.204, F.A.C., lists license categories and fees. |
| 6. Radioactive Material | Describes the type (element and mass number) of radioactive material the license authorizes for possession and use. |
| 7. Form | Describes the form of radioactive material the license authorizes for possession and use (i.e., source manufacturers and model numbers). |
| 8. Possession Limit | Lists the maximum possession limit for radioactive sealed sources. Possession of more sources than authorized by the license is a violation and may result in enforcement actions. |
| 9. Use | Describes the types of uses that are approved for the sources and devices listed in the previous items. Unauthorized use of radioactive material is a violation and may result in enforcement actions. |

License Conditions

License conditions describe requirements and limitations applicable to the radioactive materials authorized by the license.  Additional requirements and conditions may be incorporated as appropriate to protect public health and the environment.  If a licensee seeks added authorizations, supplementary license conditions may be added.  All Category 3H licenses contain conditions addressing the following:

| - Authorized location of use and storage - Enforcement provisions - Authorized User (AU) and RSO designations - Radioactive material transfer limitations - Radioactive material transportation requirements | - Enforcement provisions - Part III and IX provisions - Leak testing requirements - Inventory requirements - Licensee commitments |
| --- | --- |

II.	FILING AN APPLICATION

A.	GENERAL

Chapter 64E-5, F.A.C., this guide, forms, and other guidance documents are available on the bureau’s website: http://www.doh.state.fl.us/environment/radiation.

An application for a fixed gauge specific license must be submitted on Form DH-1054, "Application For Radioactive Materials License, Non-Human Use."  The form is included as Supplement A of this guide, and is available at http://www.doh.state.fl.us/environment/radiation/matform.htm#Forms.  Space provided on the application form is limited, so 8.5” x 11” paper should be used to append additional pages.  Each page submitted with the application should be identified and keyed to the item number on Form DH-1054 to which it applies.  Three copies of the application and all attachments must be submitted (original and two copies), with another copy retained by the applicant.  All application items must be addressed in sufficient detail to demonstrate that equipment, facilities, personnel qualifications and procedures are adequate to protect public health and safety or property.

Complete and submit the table provided as Supplement B to this guide to indicate whether model or equivalent procedures and forms have been included in the application.

| Mail to: | If using an overnight delivery service, use: |
| --- | --- |
| Florida Department of Health Bureau of Radiation Control Radioactive Materials Program 4052 Bald Cypress Way, Bin C21 Tallahassee, FL 32399-1741 | Florida Department of Health Bureau of Radiation Control Radioactive Materials Program 4042 Bald Cypress Way, Rm. 220.09 Tallahassee, FL 32399 |

With the exception of security-related information, all license applications and documents submitted to the bureau are available for review by the general public.  Do not submit proprietary information unless it is absolutely necessary for evaluation of the application.  Any request for withholding documents is subject to a determination by the department as to whether the document may actually be withheld in accordance with applicable laws and regulations.

Personal information about employees should not be submitted unless it is necessary.  Home addresses, home telephone numbers, dates of birth, and social security numbers should not be submitted unless the bureau specifically requests it.

When issued, the license will require that radioactive material be possessed and used in accordance with statements, representations and procedures provided in the application and supporting documentation (which are incorporated by referenced into the license).  Regulatory requirements specified in Chapter 64E-5, F.A.C., shall govern unless the statements, representations and procedures set forth in the license application and correspondence are more restrictive than the regulations.

B.	LICENSE FEES

The following fees are assessed:

Application fee	A non-refundable fee for processing the license application.  The amount is dependent on the category of license the applicant is seeking; refer to section 64E-5.204, F.A.C., or Regulatory Guide 6.20 for a description of application fees.  Review of the application will not begin until the proper fee is received by the department.  An application fee is also required to process an application for a new license replacing an existing license due to a change of ownership.

Annual fee	An annual fee covers department costs for administration of the materials licensing program.  The amount is dependent on the license category.  Refer to section 64E-5.204, F.A.C., or Regulatory Guide 6.20 for a description of annual fees.  Annual fees are due within 60 days of issuance of the new license; an invoice for this fee is included with the cover letter accompanying a new license.

Reclamation fee	In addition to the application and annual fees, a reclamation fee will be assessed for the Radiation Protection Trust Fund, established to pay department costs associated with a licensee’s abandonment of radioactive materials, default on lawful obligations, insolvency, or other inability to meet regulatory requirements, and to assure the protection of the public and environment.  Reclamation fees are equal to 5% of the annual fee.  Reclamation fees are due within 60 days of issuance of a new license; a fee invoice is included with the cover letter accompanying a new license.

Notes:	1.	Annual and reclamation fees are assessed on the anniversary of the license issuance date.  An invoice is sent to the licensee 30 days in advance of the due date.

2.	Fees are not assessed for license renewals, amendment requests, licensing actions, inspections initiated by the department, license terminations, or requests for regulatory information (except for document copying costs).

III.	CONTENTS OF AN APPLICATION

This section provides instructions on completing each item listed in Form DH-1054.

1.a.	NAME AND MAILING ADDRESS

List the name and mailing address of the individual or company to whom the license will be issued.  An applicant corporation or other legal entity must be listed by the legal name registered with the Department of State, Division of Corporations.  If a fictitious name is to be included, it will be identified as the name the applicant is doing business as (d/b/a) and must also be registered with the Division of Corporations.  For example, “ABC Corporation d/b/a ABC Enterprises of Florida.”  Business registration may be verified by contacting the Division of Corporations at (850) 488-9000 or on the Internet at: http://www.sunbiz.org/.

To assist in identifying the proper legal entity, applicants should also list their business’ Federal Employer Identification (FEI) or Document Number if known or applicable; the FEI or Document Number is available on the Division of Corporations website.

If the mailing address is different from where radioactive material will be used and/or stored and where license-related records will be maintained, list that address in Item 1.b.

1.b.	STREET ADDRESS AT WHICH RADIOACTIVE MATERIAL WILL BE USED and/or stored, IF DIFFERENT FROM 1.a.

Identify by street address the location where radioactive material will be used and/or stored and where license-related records will be maintained, if different that the address listed in Item 1.a.  Non-contiguous locations of use and storage may require a separate license; refer to section 64E-5.213, F.A.C.

2.a.	LICENSE FEE CATEGORY

Indicate the appropriate license fee category; for fixed nuclear gauges, list category 3H.  Refer to section 64E-5.204, F.A.C., or to Regulatory Guide 6.20 for a listing of fees

2.b.	LICENSE FEE ENCLOSED

Indicate the amount of the enclosed license application fee in the space that is provided.  Refer to section II.B. of this guide for a description of licensing fees.

3.	THIS IS AN APPLICATION FOR:

Mark the appropriate choice; if submitting an amendment request or a renewal application, indicate the applicable radioactive materials license number in the space provided.

4.	INDIVIDUAL USERS

List each individual to be designated as an authorized user (AU) of radioactive material (i.e., a gauge operator).  A minimum of two AUs is required unless otherwise approved.  Two types of AUs may be approved under a fixed gauge license, as described below.

- Basic AUs are individuals authorized to perform “basic services”: routine operations such as shutter checks, maintenance, physical inspections/inventories, and leak tests.  Such workers must complete at least 8 hours of training covering the subjects listed in subsection 64E-5.1307(1), F.A.C.

- Advanced AUs are individuals authorized to perform “advanced services”: non-routine operations such as gauge installations, relocations, radiation surveys and limited maintenance and repairs.  Such workers must complete at least 40 hours of training covering the subjects listed in subsection 64E-5.1313(2), F.A.C.

Item 4 should identify each listed individual user as either a basic AU or an advanced AU.

5.	RADIATION SAFETY OFFICER (RSO)

Provide the name of the individual assigned the position of RSO.  This person is designated by and responsible to management for implementation of the radiation safety program and for ensuring compliance with the applicable regulations and license provisions.  As a minimum, the RSO must have sufficient training and experience to be a basic authorized user of the requested radioactive materials.  Additional training in administration of a radiation protection program is recommended for the RSO position.

6.	TRAINING AND EXPERIENCE IN RADIATION SAFETY

a.	FORMAL TRAINING IN RADIATION SAFETY

Submit documentation of radiation safety training for each individual listed in Items 4 and 5 of the application, demonstrating that the workers meet the training requirements for basic or advanced AUs, as applicable.  Restrict training documentation to relevant information; i.e., demonstrating that the individual has the radiation safety training and experience specific to the requested activities to be conducted.  Appropriate training certificates such as those provided by gauge manufacturers or other approved third parties are acceptable.  Such certificates may need to be supplemented with documentation of completion of training in the applicant’s operating and emergency (O&E) procedures to satisfy Chapter 64E-5, F.A.C., Part XIII training requirements, because third party trainers may not provide such training.

b.	EXPERIENCE

Describe any additional relevant work experience with radiation and where the experience was obtained.  Descriptions of experience are typically unnecessary unless seeking approval to act as an instructor for in-house radiation safety training.  Do not include individuals’ birth dates or social security numbers.

7.	RADIOACTIVE MATERIAL

a.	ELEMENT AND MASS NUMBER

List each type of radioactive material requested; refer to the example provided below.

b.	CHEMICAL AND/OR PHYSICAL FORM

Complete for each type of radioactive material requested.  State the name of the source manufacturer and the source model number; refer to the example provided below.

c.	MAXIMUM AMOUNT TO BE POSSESSED AT ANY ONE TIME

Complete for each radioactive material requested.  Indicate the total number of sources and maximum activity per source.

Example:

| (a) ELEMENT AND MASS NUMBER | (b) CHEMICAL AND/OR PHYSICAL FORM | (c) MAXIMUM AMOUNT TO BE POSSESSED AT ANY ONE TIME |
| --- | --- | --- |
| 1. Cesium 137 | Sealed source (XYZ, Inc. Model 123) | 2 sources; not to exceed 100 millicuries each |

If authorization for generally licensed (GL) fixed gauges or other GL devices is sought, include a request for GL sources and devices.

Example:

| (a) ELEMENT AND MASS NUMBER | (b) CHEMICAL AND/OR PHYSICAL FORM | (c) MAXIMUM AMOUNT TO BE POSSESSED AT ANY ONE TIME |
| --- | --- | --- |
| 2. Radioactive material Sealed sources distributed to a general licensee per 64E-5.206(1) & (4), F.A.C. |  | No single source to exceed that quantity authorized for the general license |

Note:	Applicants seeking authorization for GL sources and devices must comply with section 64E-5.1308, F.A.C., which describes requirements for GL devices possessed under a specific license.  Annual inspections/inventories must include the GL sources, and the sources must be leak tested at the interval specified by the manufacturer.

8.	PURPOSE FOR WHICH RADIOACTIVE MATERIALS WILL BE USED

Complete for each radioactive material requested.  Include the name of the manufacturer and model of the device or source holder in which each source is used or stored.

Example:

1.	For use in XYZ Corporation Model 2000 source holders contained in XYZ Model 2003 gauges for level measurements.

If authorization for GL gauges or other GL sources/devices is requested, describe the intended use of the GL sources and devices.

Example:

2.	To be used in devices approved for receipt under general license provisions.

CURRENT INVENTORY

Applicants for renewal of an existing license must include an inventory of all sealed sources and devices currently possessed.  List all generally licensed and specifically licensed sources and devices, and indicate the licensing designations for each (i.e., GL or SL).  If in possession of any exempt sources (e.g., check sources), it is recommended that they be included in the inventory in order to avoid any confusion over their licensing status.  If exempt sources are included, identify their exempt status on the inventory form.

9.	RADIATION DETECTION INSTRUMENTS

Possession of a radiation detection instrument (survey meter) is recommended for all gauge licensees, but is not required for applicants seeking authorization to perform basic (routine) gauge-related activities.  For such applicants, this item may be marked “not applicable” or  “N/A.”  Applicants selecting this option must describe in their procedures how a suitable survey meter will be available for use in identifying and assessing radiological hazards.  Access to a survey meter may be obtained by making arrangements with a local emergency responder (hazmat team, etc.) or with another licensee, such as a fixed gauge licensee, a testing lab, a hospital or outpatient clinic, a radiation safety consulting company, etc.

Applicants seeking authorization to perform advanced (non-routine) activities on fixed gauges are required to have at least one survey meter available with range detection capabilities suited to the gauge(s) possessed.

There are other situations where a survey meter is needed to determine whether the integrity of a gauge’s shielding or source has been compromised (e.g., receipt of a damaged gauge or damage to a gauge as a result of an industrial accident).  Such incidents necessitate seeking technical assistance to arrange for a timely evaluation of the gauge.  Therefore, emergency procedures need to include instructions regarding access to a survey meter (Appendix J to this guide is a model emergency procedure).

10.	CALIBRATION OF INSTRUMENTS

If radiation detection instruments will be used, mark the appropriate box to indicate who will perform calibrations.

a.	CALIBRATED BY SERVICE COMPANY

If a service company will be used, list the vendor(s) name, address, license number and the government agency that issued the company’s license (i.e., NRC or a state agency).  Survey instruments must be calibrated at least annually per subsection 64E-5.314(2), F.A.C.

b.	CALIBRATED BY APPLICANT

If seeking approval to calibrate instruments in-house, submit detailed information describing the facilities, equipment, personnel, and procedures to be used to perform the calibrations.  Contact the BRC for additional guidance on equipment calibration requirements.  Note: in-house calibration requires use of reference sources; list each requested calibration/reference source in Item 7.

11.	PERSONNEL MONITORING DEVICES

If conducting personnel monitoring (PM), complete Items a., b., c. and d. and address PM procedures in the radiation protection program.  If no PM program will be implemented, write “N/A.”

Unless otherwise authorized, subsection 64E-5.1310(2), F.A.C., requires individuals performing advanced services on fixed gauges to wear a whole body PM badge.  Applicants performing basic services only may elect not to conduct PM if able to demonstrate that workers are unlikely to exceed 500 millirem/year from gauging operations.  Appendix G is a model procedure for documenting a decision not to perform PM.

Common PM badges include film badges, thermoluminescent dosimeters (TLDs) and optically stimulated luminescent dosimeters (OSLDs), which are described below.  PM badges must be capable of detecting the type of radiation (e.g., beta, gamma, neutron) emitted by the gauges possessed by the licensee.

Badge processors must hold accreditation form the National Voluntary Laboratory Accreditation Program (NVLAP) of the National Institute of Standards and Technology.  A list of NVLAP-accredited badge vendors is available at http://ts.nist.gov/ts/htdocs/210/214/scopes/dosim.htm.  A list of commercial PM badge vendors is available from the bureau upon request.

Each order of badges includes a control badge for measuring the amount of background radiation the badges receive each monitoring period.  This enables background radiation to be subtracted from the total reading to provide an accurate record of each worker’s occupational exposure.  When not in use, PM badges should be stored with the control badge in an environment protected from radiation, chemicals, excessive heat, light, moisture, etc. to ensure accurate dose records.  The control badge must be returned with the other PM badges each monitoring period.

Film badges are small pieces of x-ray film contained in a plastic holder.  The film darkens in proportion to the amount of radiation it has been exposed to, so the film’s optical density provides a measurement of the wearer’s radiation exposure.  Film badges must be exchanged monthly.

TLDs are PM badges that contain small crystals capable of storing some of the energy from radiation.  Heating the crystals releases the stored energy as light.  The amount of light released is proportional to the amount of radiation the TLD badge received, which is measured to determine the badge wearer’s dose.  TLDs must be exchanged at least every three months.

OSLDs measure radiation using a thin layer of aluminum oxide.  A laser light stimulates the aluminum oxide after use, causing it to become luminescent in proportion to the amount of radiation exposure.  OSLDs must be exchanged at least every three months.

12.	FACILITIES AND EQUIPMENT

Facilities and equipment must be adequate to protect health and minimize danger to life and property.  Fixed gauges must be secured to prevent unauthorized access or removal. (e.g., permanently mounted, located in a locked room, or chained and locked to a storage rack).

Fixed gauges may be located in harsh environments resulting from  pressure, vibration, mounting height/method, temperature, humidity, air quality, corrosive atmospheres or chemicals, possible impact or puncture conditions, and fire, explosion and flooding potentials.  Applicants should consult the gauge manufacturer’s literature as well as the device’s sealed source and device (SSD) certificate to determine their gauges’ limitations and/or other considerations of use.  If a proposed use conflicts with imposed by the manufacturer and/or SSD certificate, specific information must be provided to demonstrate that the proposed conditions of use will not impact the safety or integrity of the device.  The National Sealed Source and Device Registry is available on the Internet at http://www.hsrd.ornl.gov/nrc/sources/index.cfm.

Submit an annotated diagram of the facility identifying all areas where gauges will be installed, and where gauges are stored prior to installation or awaiting shipment for disposal or repair.  In addition, identify all occupied workstations and areas adjacent to installed or stored gauges.

Describe the security measures in place to prevent unauthorized access or removal of radioactive materials.  Describe the area where gauges in storage are secured and the security measures in place to prevent unauthorized access.  Refer to section 64E-5.320, F.A.C., for regulations on security of radiation sources.

Describe any anticipated environmental conditions at the location of use that have the potential to adversely affect gauge operation and visibility (e.g., excessive temperature, vibration, caustic chemicals, airborne particulates), and describe what protective measures will be used to minimize the impact of the adverse conditions.  Examples of mitigation methods include use of protective barriers and increased frequency of routine gauge inspections.

Describe any auxiliary shielding or barriers installed around source housings or gauging systems to limit access and/or reduce scattered radiation in occupied areas.

13.	RADIATION PROTECTION PROGRAM

Submit a detailed description of the proposed radiation protection program, which must include the components described below.  The appendices and exhibits included with this guide are model procedures and forms that may be adopted by including them as part of the submitted radiation protection program, or used as guides for developing equivalent procedures and forms.

Reminder: Complete Tables 1 – 3 of Supplement B of this guide to indicate whether model or equivalent procedures and forms have been submitted, and attach a copy to Form DH-1054.

Member of the Public (MOP) Dose Limit Compliance Study	Appendix A

Applicants for a new license must submit proposed procedures for demonstrating compliance with the public dose limits specified in section 64E-5.313, F.A.C. (2 millirem in any one hour and 100 millirem per year).  Applicants for renewal of an existing license must submit a completed study demonstrating compliance with the limits.  Appendix A is a model study.

ALARA Policy	Appendix B

Applicants must submit a policy describing management's commitment to the ALARA philosophy of maintaining doses as low as reasonably achievable, and a description of the commitments of management and workers for implementing the policy.  Appendix B is a model ALARA policy.

Duties and Responsibilities of the Radiation Safety Officer (RSO)	Appendix C

Applicants must submit a description of the RSO’s duties and responsibilities that includes the duties listed in section 64E-5.1305, F.A.C.  Appendix C is a model procedure.

Radiation Safety Training Program	Appendix D

Multiple training requirements are associated with fixed gauge operations, as described below.

- Radiation awareness training (“Instructions to Workers”) must be provided to personnel engaged in licensed activities (AUs and workers under their supervision).  This training is also recommended for individuals that work in the vicinity of gauges.  Section 64E-5.902, F.A.C., specifies radiation awareness training requirements.  The rule does not specify the minimum duration for this training, because the amount of training needed will vary depending on the scope of the radiological hazards present in the applicant’s workplace.

- Hazmat employee training must be provided to any worker with job functions associated with gauge shipments.  Hazmat employee training is specified in 49 CFR Part 172, Subpart H.  U.S. DOT regulations are incorporated by reference in Chapter 64E-5, F.A.C.  The rule does not specify the minimum duration for this training.

- Basic AU training must be provided to workers independently performing basic (routine) gauge operations or supervising such activities by other workers.  Basic AU training requirements are specified in subsection 64E-5.1313(1), F.A.C.  Basic AUs must complete a minimum of 8 hours training covering the subjects listed in section 64E-5.1307, F.A.C.

- Advanced AU training must be provided to workers independently performing advanced gauge operations or supervising such activities by other workers.  Advanced AUs must complete at least 40 hours of training that covers the subjects listed in subsection 64E-5.1313(2), F.A.C.

The training program must describe how the above training requirements will be addressed.  Because the topics that must be addressed to satisfy radiation awareness and hazmat employee training overlap with the topics that must be covered during fixed gauge AU training, training requirements may be addressed concurrently.  Training can be provided by qualified third parties, in-house, or by using a combination of the two.  Appendix D is a model radiation safety training program aimed at licensees using an approved third party training course, supplemented by in-house training in the licensee’s O&E procedures.

Applicants seeking to conduct in-house training must address the requirements specified in subsection 64E-5.1307(2), F.A.C., by submitting a detailed description of the training program for review and approval, including a description of training reference materials and a sample exam.

Operating and Emergency Procedures

Sections 64E-5.208 and 64E-5.1302, F.A.C., require establishment and implementation of O&E procedures that provide instructions adequate to ensure safety to workers, the public and to property.  As a minimum, O&E procedures must include the procedures described below.  Applicants seeking approval to perform advanced (non-routine) services must submit procedures addressing those activities.

- Operating Procedure	Appendix E

Due to the wide variety of gauges and gauging applications, it is not feasible to provide a model operating procedure that addresses every type of gauge and application.  Applicants must refer to their gauge manufacturer operation/maintenance manuals to ensure that all necessary instructions are included.

Appendix E provides generic instructions on availability of personnel and procedures, general rules of use/ALARA principles, radiation surveys, security, and basic and advanced gauge services.  Operating procedures must include instructions on those topics; additional guidance is provided below.

Availability of Personnel and Procedures

Section 64E-5.1313, F.A.C., requires an authorized user (basic or advanced AU) to be available at all times when fixed gauges are in use.  "Available" means that as long as fixed gauges are on site (even if they are locked out), an AU must either be on-site or able to be on-site within an hour's notice.

Operating procedures must include a commitment that AUs will have access to the licensee’s O&E procedures, as well as the manufacturer operation/maintenance manual for each type of gauge possessed.

General Rules of Use/ALARA Principles

Instructions on use of dose minimization techniques should be supplemented by any device-specific instructions provided by the gauge manufacturer.

Radiation Surveys

If damage to a gauge is suspected, a radiation detection instrument must be used to measure the radiation levels.  If a survey meter is not on site, the RSO must be able to make arrangements for the gauge to be surveyed.  Exhibit B is a generic form illustrating how radiation surveys performed on density gauges installed on piping may be documented.  Applicants should contact their gauge manufacturer(s) to obtain model-specific survey report forms that are applicable to the gauges and configurations in use at their facility, and use those forms to document radiation surveys.

Security

Gauges must be stored and installed in a manner that secures them from unauthorized access or removal.  Additional controls (e.g., cages, fencing, guards, surveillance monitoring systems) should be utilized as appropriate to enhance gauge security.

Basic and Advanced Services

Examples of basic gauge services include shutter checks, leak tests, physical inspections and inventories, and normal maintenance (e.g., removal of dirt and debris, rust removal, painting).  Only AUs or workers under their supervision are authorized to perform basic gauge services, which must be performed in accordance with manufacturer instructions.

Advanced activities include gauge installations, advanced maintenance or service, relocations, and removal from service.  Non-routine gauge maintenance or repair that requires removal of the source is prohibited.  Only advanced AUs or workers under their direct supervision and in their physical presence are authorized to perform advanced activities.  Assigned PM badges must be worn and a survey meter must be used when performing advanced services.  Survey results must be documented and maintained.  Manufacturer instructions must be strictly followed.

- Personnel Monitoring Procedure	Appendix F

If PM is conducted, a procedure is required to provide instructions on proper use, exchanges, use of spare badges, lost or damaged badges, and PM record-keeping requirements.  Appendix F is a model PM procedure.  Exhibits E, F and G are model forms for addressing declared pregnant female requirements.

If personnel monitoring will not be conducted, attach documentation that includes measurements and/or calculations demonstrating that personnel working with fixed nuclear gauges are not likely to exceed 500 millirem per year.

- Posting and Labeling Procedure	Appendix H

The posting/labeling procedure must address requirements for labeling gauges, posting areas where gauges are installed or stored, and posting documents specified in section 64E-5.901, F.A.C.  Appendix H is a model posting/labeling procedure.

When appropriate instrumentation (a survey meter) is available, a radiation survey of ambient radiation levels around all gauges should be performed to verify compliance with area posting requirements.  When such instrumentation is unavailable, information on gauge radiation levels provided by manufacturers should be used to ensure compliance.

- Lock-out Procedure	Appendix I

The procedure must provide instructions for when gauge lock-outs are required and how they are performed, and must comply with U.S. Occupational Safety and Health Administration standards for the control of hazardous energy, as specified in Title 29, Code of Federal Regulations, Part 1910, section 147.  Appendix I is a model lock-out procedure.

- Emergency Procedures	Appendix J

The procedures must provide instructions for responding to the loss, theft or damage of a gauge, and must include emergency notification numbers for the RSO and bureau.  Appendix J is a model emergency procedure.

- Procedures for Ordering, Receiving, Opening and Shipping	Appendix K

Packages Containing Radioactive Material	Exhibits C and D

Procedures must address preparation and handling of incoming and outgoing shipments of radioactive material transported by common carriers, and if applicable, for gauges transported as private use shipments.  The instructions must conform to current U.S. DOT regulations.  Sample shipping papers and emergency response information must be provided.  Appendix K is a model procedure, Exhibit C is a sample shipping paper for private use shipments, and Exhibit D is a model emergency response information sheet.

- Leak Testing Procedure	Appendix L

Sealed sources in gauges must be tested at regular intervals to ensure that the radioactive material is not leaking contamination.  Leak test (LT) requirements are specified in section 64E-5.1303, F.A.C.  The LT procedure must either specify the manufacturer name and model number of the LT kit used and the name of the vendor contracted to perform LT sample analysis, or provide a commitment to use licensed vendors and approved LT kits.

The procedure must specify the leak test interval and include instructions on collecting LT samples.  Such information is available from gauge manufacturers.  Appendix L is a model LT procedure.

If authorization to perform in-house leak test analysis is sought, additional procedures addressing sample analysis must be submitted.  The procedure must describe the instrumentation used to perform the analysis and include step-by-step instructions for calculating instrument efficiency and performing sample counts.  Additional guidance on in-house leak test analysis is available from the bureau.

- Inventory Procedure	Appendix M

The procedure must provide instructions for performing annual physical inventories and inspections of all generally and specifically licensed (GL and SL) sources.  Inspections must evaluate the physical condition of each gauge and the gauge labels, posted warning signs, and the ambient environmental conditions impacting the gauges.  A sample inventory form is also required.  Appendix M is a model inventory procedure; Exhibit A is a model inventory form.

- Record Retention Procedure	Appendix N

Certain records must be retained for specified periods of time for compliance purposes.  These intervals have been established for BRC inspection staff and other authorized entities (e.g., U.S. DOT) to have access to the documents as required by the regulations.  Appendix N provides a model procedure addressing record retention requirements applicable to fixed gauge licensees.

- Notification and Reporting Procedure	Appendix O

Notification and reporting requirements are specified in Parts II and III of Chapter 64E-5, F.A.C.  Appendix O provides a model procedure summarizing notification and reporting requirements applicable to fixed gauge licensees.

- Gauge Installation/Relocation/Removal Procedure

A model procedure for gauge installations, relocations and removals is not included in this guide, because the instruction can vary significantly from model to model.  Applicants seeking approval to conduct such activities must contact their gauge suppliers for guidance on developing procedures for these tasks.

14.	WASTE DISPOSAL	Appendix P

Submit a procedure describing how radioactive sources contained in fixed gauges will be disposed.  The procedure must include a commitment that sealed sources will be disposed of either by returning it to the manufacturer or by transferring it to a specifically licensed recipient.  Indefinite storage is not a recommended disposal option.  Appendix P is a model waste transfer/disposal procedure.  Please note that low-level radioactive waste brokers and most gauge manufacturers require a fee to accept gauge sources for disposal.

15.	CERTIFICATE

A radioactive materials license is a legal document.  License applications and license-related correspondence must be signed and dated by an individual (certifying official) authorized to make legally binding statements for the applicant.  Examples of positions that are recognized as certifying officials include owner, president, vice president, chief executive officer, chief operating officer, etc.  Positions that are not recognized as certifying officials include RSO, environmental health & safety director, and plant manager.

A certifying official may delegate authority to make legally binding statements to specific individuals or positions (e.g., manager, director, RSO) by submitting a written statement authorizing the delegation.  Exhibit H is a model form for documenting a delegation of authority.

16.	* NEW - INCREASED SECURITY CONTROLS – IF APPLICABLE

The U.S. Nuclear Regulatory Commission (NRC) and its Agreement States have implemented increased controls for licensees that possess certain radioactive materials in quantities of concern.  The NRC has determined that additional requirements need to be implemented to supplement existing regulatory requirements in 10 CFR Sections 20.1801-20.1802.  The licensee shall comply with the requirements described in the document titled Bureau of Radiation Control – Increased Security Controls Requirements - October 14, 2005, and its attachments, which are incorporated by reference, titled "Increased Controls for Licensees that Possess Sources Containing Radioactive Material Quantities of Concern IF:

The total sources or devices that you wish to posses meet or exceed the activities listed below.  Be advised that until proved otherwise all sources will be considered co-located and must be added together.  Co-located is defined as breaching a common physical security barrier to allow unrestricted access to sources or devices.  In other words multiple fixed gauges in a facility where all gauges are accessible after passing through a perimeter security check point, or by breaching a perimeter fence, would be considered co-located. However, if additional physical security barriers are present within the facility that would prevent access to quantities of radioactive material exceeding Table 1 quantity listed below, then the sources or devices are not considered aggregated or co-located and implementation of the Increased Controls (ICs) is not required. This can be accomplished by, but is not limited to, locked enclosures, tamper proof mounting bolts (one way threading), robust cables, chains with locks.  Therefore, in order for a license to consider a physical barrier to be effective, the licensee must ensure the barrier cannot be bypassed or easily defeated by using commonly available tools.

| Radionuclide | Quantity of Concern1 (TBq) | Quantity of Concern2 (Ci ) |  | Radionuclide | Quantity of Concern1 (TBq) | Quantity of Concern2 (Ci ) |
| --- | --- | --- | --- | --- | --- | --- |
| Am-241 or Am-241:Be | 0.6 | 16 |  | Pm-147 | 400 | 11,000 |
| Cf-252 | 0.2 | 5.4 |  | Pu-238 | 0.6 | 16 |
| Cm-244 | 0.5 | 14 |  | Pu-239:Be | 0.6 | 16 |
| Co-60 | 0.3 | 8.1 |  | Se-75 | 2 | 54 |
| Cs-137 | 1 | 27 |  | Sr-90 (Y-90) | 10 | 270 |
| Gd-153 | 10 | 270 |  | Tm-170 | 200 | 5,400 |
| Ir-192 | 0.8 | 22 |  | Yb-169 | 3 | 81 |

1 The aggregate activity of multiple, collocated sources of the same radionuclide should be included when the total activity equals or exceeds the quantity of concern.

2 The primary values used for compliance with this Order are TBq.  The curie (Ci) values are rounded to two significant figures for informational purposes only.

3 Radioactive materials are to be considered aggregated or collocated if breaching a common physical security barrier (e.g., a locked door at the entrance to a storage room) would allow access to the radioactive material or devices containing the radioactive material.

4 If several radionuclides are aggregated, the sum of the ratios of the activity of each source, i of radionuclide, n, A(i,n), to the quantity of concern for radionuclide n, Q(n), listed for that radionuclide equals or exceeds one.   [(aggregated source activity for radionuclide A) ÷ (quantity of concern for radionuclide A)] + [(aggregated source activity for radionuclide B) ÷ (quantity of concern for radionuclide B)] + etc. >1

IV.	LICENSE AMENDMENTS

Licensees are required to conduct operations in accordance with applicable regulations and the statements, representations and procedures contained in the license application and supporting documents.  The license must be amended if any changes are planned.  Submittal of an amendment request does not allow immediate implementation of proposed changes.  Until the license has been amended to reflect approval of the change(s), the licensee must comply with the original terms and conditions of the license.

Applications for license amendments may be submitted in letter form or on Form DH-1054, “Application For Radioactive Materials License, Non-Human Use.”  The request must be dated and signed by a certifying official, identify the license by name and number, be submitted in triplicate, and clearly describe the nature of the changes, additions or deletions requested.  References to previously submitted documents must be specific and identify the applicable information by date, page and paragraph.  The licensee must maintain a copy of the submitted and referenced documentation on file for inspection.  To prevent the potential for identity theft, do not submit documentation that lists individuals’ social security numbers or birth dates.

V.	LICENSE RENEWAL

Absent any actions by the department or the licensee, a license remains in effect for five years.  An application for license renewal must be received by the department at least 30 days prior to the expiration date.  This filing will ensure that the license does not expire until final action on the application has been taken, as provided for by subsection 64E-5.207(3), F.A.C.  If the application is received less than 30 days before the expiration date, the facility or individual may be without a valid license when the license expires.  Renewal applications should be filed using Form DH-1054, “Application For Radioactive Materials License, Non-Human Use” (Supplement A of this guide).  Renewals require submittal of an entire new application, completed as if it were an application for a new license, with complete and up-to-date information about the applicant's radiation protection program, demonstrating compliance with all licensing and regulatory requirements in effect at the time of renewal.  Renewal applications should be submitted without reference to documentation and information submitted previously.  To prevent the potential for identity theft, never submit documentation that lists individuals’ social security numbers or birth dates.

VI.	LICENSE TERMINATION

Prior to license termination, the licensee must properly dispose of all licensed radioactive material possessed.  Complete Form DH-1059, “Certificate – Disposition of Radioactive Material” to satisfy the requirements of section 64E-5.214, F.A.C., and submit it to the Bureau before the expiration date of the license with a request that the license be terminated.

1.20 Intro  2/14/28