I.	Scope and Purpose

This procedure establishes the minimum requirements for locking out fixed nuclear gauges when maintenance or servicing is performed on or near gauges such that workers could be exposed to the gauge’s primary radiation beam or scattered radiation.  Compliance with this procedure is required to ensure that gauges are properly locked and/or tagged out before personnel perform any work where installed gauges could cause unnecessary radiation exposures.

As used in this procedure, “lock-out/tag-out” refers to methods used to safeguard workers from exposure to ionizing radiation emitted by radioactive sources contained in fixed gauges installed on process or other equipment.  Lock-out devices provide protection by serving as positive restraints that no one can remove without a key or other unlocking mechanism, or through extraordinary means, such as bolt cutters.  Tag-out devices, by contrast, are prominent warning devices used to warn workers not to open a gauge shutter or otherwise expose a gauge source while the service or maintenance activity is being performed.  Tag-out devices are easier to remove and, by themselves, provide workers with less protection than do lock-out devices.

All workers are required to comply with the restrictions and limitations imposed upon them when conditions require gauge lock-out/tag-out.

II.	Conditions Requiring Lock-Out

A gauge source holder will be locked-out by locking the on/off or shutter mechanism into a safe position – the “off” or closed position:

- Prior to any work being performed in the immediate vicinity of a gauge radiation beam when a distance or gap exists between a gauge’s radioactive source and the radiation detector that permits entry of all or a portion of a person’s body into the primary radiation beam;

- During any manipulation of a gauge, including the source holder or the detector, which involves physical movement of the device or separation from a pipe, vessel, etc. including installation, relocation or storage;

- When individuals are working on or adjacent to a gauge during periods of shutdown;

- Whenever an individual enters a vessel in which such a gauge is located; and

- Whenever a vessel with such a gauge is empty and an individual is working around the exterior of the vessel.

III.	Lock-out/Tag-out Specifications

Lock-out and tag-out devices will be singularly identified, will be the only devices used for lock-outs/tag-outs, will not be used for other purposes, and will satisfy the specifications described below.

Lock-out devices will consist of either a key or combination lock capable of holding the gauge in the safe (closed) position such that the gauge cannot operate until the lock-out device is removed.  Lock-out devices will be substantial enough to prevent removal without the use of excessive force or unusual techniques.

III.	Lock-out/Tag-out Specifications    (contd.)

Tag-out devices will consist of a durable tag and a means of attachment that can be securely fastened to the gauge to indicate that the gauge may not be operated until the tag-out device is removed.  Tag-out devices will be substantial enough to prevent inadvertent or accidental removal, and able to withstand the ambient environment for the maximum period of time that exposure is expected.  Tag-out device attachments will be of the non-reusable type, attachable by hand, self-locking, and non-releasable with a minimum unlocking strength of no less than 50 pounds, with the general design and basic characteristics at least equivalent to a one-piece, all-environment-tolerant nylon cable tie.  Tag-out devices will warn against hazardous conditions if the gauge is operated and must include a legend such as Do Not Open or Do Not Operate.  Tags shall be legible and understandable to all personnel who may be in the area.

Lock-out and/or tag-out devices will indicate the identity of the individual applying the device(s).  Lock-out and/or tag-out devices will be standardized in at least one of the following criteria: color; shape; or size, and the print and format of tag-out devices will be standardized.

IV.	Lock-out/Tag-out Sequence

Only the Radiation Safety Officer (RSO) or RSO designees are authorized to lock-out/tag-out a gauge.  All workers, upon observing a gauge that is locked and tagged, shall not attempt to operate the gauge or remove the lock and tag.

A.	When work is required on or near a gauge, notify all affected personnel that the gauge shutter must be closed, locked-out, and tagged prior to initiating the work.

B.	The RSO or another authorized fixed gauge user (AU) will lock and tag the gauge in accordance with manufacturer recommendations, using lock-out/tag-out devices meeting the specifications described in this procedure.

C.	When locking out a gauge, the on/off or shutter mechanism will be tagged to indicate that the gauge is locked out.  If a gauge is incapable of being locked out, a tag-out device must still be used.

D.	The RSO or RSO designee will verify that the gauge has been effectively locked out.  Verification may include a physical inspection of the shutter mechanism and/or ON/OFF indicator, an inspection of the data display, and performance of a radiation survey (if a survey meter is available).

E.	A warning sign will be posted at each entryway to areas where it is possible to be exposed to the primary radiation beam from the gauge.  Such warning signs will include safety instructions (e.g., “Contact the Radiation Safety Officer Before Entering Vessel”).