Sealed sources contained in fixed gauges must be tested at regular intervals to ensure that the radioactive material is secure within its capsule and not leaking contamination.  Leak test (LT) requirements are specified in section 64E-5.1303, Florida Administrative Code.

I.	Leak Test Frequency

The frequency for testing a gauge for leakage can vary from model to model; typical intervals are 1 or 3 years.  Listed below are the gauge models currently in use, and their required test intervals.

| - Gauge models that will be leak tested at least every 12 months: |  |
| --- | --- |
| - Gauge models that will be leak tested at least every 36 months: |  |
| - Gauge models that will be leak tested at least every ___ months: |  |

II.	Leak Test Kit

Only LT kits provided by licensed LT vendors will be used to sample (smear) sealed sources contained in fixed gauges.

III.	Taking the Leak Test Sample

LT samples will be taken by an Authorized User (AU) or other trained individuals authorized to perform the procedure.  If the AU has been assigned a personnel monitoring badge, the assigned badge will be worn when collecting the LT sample.  Samples will be taken in accordance with the written instructions provided by the supplier of the LT kit and/or the gauge manufacturer.  Leak test samples should be collected at the most accessible area where contamination would accumulate if the sealed source were leaking.  If available, use a survey meter to monitor radiation levels.  Prepare a separate wipe sample (e.g., cotton swab or filter paper) for each source.  For each source tested, list identifying information (gauge model and serial number, radionuclide, activity, etc.), and number each wipe to correlate the LT sample to the source being tested.

IV.	Leak Test Sample Analysis

Analysis of LT samples will be performed only by vendors specifically licensed to provide the service by the Florida Bureau of Radiation Control (BRC), the U.S. Nuclear Regulatory Commission, or other state radiation control agencies.

V.	Leak Test Records

If a test indicates a gauge’s sealed source is contaminated, the gauge will be removed from service and the BRC will be notified immediately (407/297-2095).  A written report on the leaking source will be submitted to the BRC within 5 days.  The report will describe the equipment involved, the test results, and the corrective actions taken (i.e., gauge removed from service until repaired; radiation surveys conducted to determine presence of contamination; decontamination as necessary).

Leak test records will be retained for 3 years for inspection purposes.  The records will include the following information:

- Each source’s manufacturer name, model, and serial number;

- The identity of each sealed source radionuclide and its estimated activity, expressed in millicuries (or becquerels);

- The measured activity of each leak test sample, in microcuries (or Bq);

- The date the sample was collected; and

- The signature of the Radiation Safety Officer (or the RSO’s designee).