Regulatory Guide 1.50
Revised:  July 2013
Revision 1

Instructions for Preparing Applications
for Radioactive Materials Licenses Authorizing

UNSEALED AND/OR SEALED SOURCES

(For In-vitro/clinical laboratories, academic, research, NUCLEAR SERVICE and other miscellaneous uses subject to PART xiii OF CHAPTER 64e-5, fLORIA ADMINISTRATIVE CODE)

TABLE OF CONTENTS

Page

I.	INTRODUCTION	1

A.	Purpose of Guide	1

B.	Applicable Regulations	2

C.	License Requirements and Restrictions	3

II.	FILING AN APPLICATION	4

A.	General	4

B.	License Fees	5

III.	CONTENTS OF AN APPLICATION	6

Item 1.a.	Name and Mailing Address of Applicant	6

Item 1.b.	Street Address(es)	5

Item 2.a.	License Fee Category	5

Item 2.b.	License Fee Enclosed	6

Item 3.	Application Type	6

Item 4.	Individual Users	7

Item 5.	Radiation Safety Officer	7

Item 6.	Training and Experience in Radiation Safety	7

Item 7.	Radioactive Material	8

Item 8.	Use of Radioactive Material	9

Item 9.	Radiation Detection Instruments	9

Item 10.	Instrument Calibration	9

Item 11.	Personnel Monitoring Devices	10

Item 12.	Facilities and Equipment	11

Item 13.	Radiation Protection Program	11

Item 14.	Waste Disposal	15

Item 15.	Certificate	15

IV.	LICENSE AMENDMENTS	16

V.	LICENSE RENEWAL	16

VI.	LICENSE TERMINATION	16

i

TABLE OF CONTENTS

APPENDICES

App. A.	Member of the Public (MOP) Dose Limit Compliance Study

App. B.	ALARA Policy

App. C.	Duties and Responsibilities of the Radiation Safety Officer

App. D.	Radiation Safety Training Program

App. E.	Personnel Monitoring Procedures

App. F.	Operating Procedures

App. G.	Procedures for Ordering, Receiving, Opening and Shipping Packages Containing Radioactive Material

App. H.	General Rules of Use of Unsealed Sources

App. I.	Contamination Control

App. J.	In-Vivo Thyroid Bioassay

App. K.	In-Vitro Urinalysis Bioassay

App. L.	Emergency Procedures

App. M.	Leak Testing Frequency, Collection and Records

App. N.	Leak Test Analysis Procedure

App. O.	Inventory Procedure

App. P.	Record Retention Procedure

App. Q.	Notification and Reporting Procedure

App. R.	Radiation Detection Instrument Calibration Procedure

App. S.	Transfer/Disposal Procedure

EXHIBITS

Ex. A	Inventory Form

Ex. B.	Shipping Paper – Private Use Shipment

Ex. C.	Emergency Response Information

Ex. D.	Instructions for Women Working With Radiation

Ex. E.	Instructions to Declared Pregnant Women

Ex. F.	Declaration of Pregnancy

Ex. G.	Delegation of Authority to Make Legally Binding Statements

SUPPLEMENTS

Supp. A.	Form DH-1054 “Application for Radioactive Materials License – Non-Human Use”

Supp. B.	List of Attachments Provided With Application

Supp. C.	Unsealed/Sealed Sources License Application Checklist

ii

I.	INTRODUCTION

A.	PURPOSE OF GUIDE

This guide provides instructions for preparing an application for a state of Florida radioactive materials license authorizing possession and use of unsealed and/or sealed sources in varying settings ranging from in-vitro, clinical and environmental laboratories, academic, research and development, nuclear service and other miscellaneous uses subject to Part XIII of Chapter 64E-5, F.A.C.  Due to the varying nature and scope of licensed operations for which this regulatory guide was developed, the applicability of the procedures contained in this guide will vary for each applicant.  In turn, additional information or procedures may be requested when necessary to ensure that an adequate radiation protection program has been established or to address unique aspects of the proposed use of radioactive materials.

License Type:  General and Specific

There are two license types: general license and specific license.  Certain concentrations, quantities, and items are exempt from the regulatory requirements and may be received without a general or specific license.  For additional information concerning exempt concentrations, quantities, and items, review section 64E-5.303, F.A.C.

A general license does not require the filing of applications with the department or the issuance of licensing documents, although the filing of a certificate with the department is required of certain users as specified in subsections 64E-5.206(7) and (8), F.A.C.  The license becomes effective upon receipt of the radioactive material.  The supplier is required to notify Florida whenever radioactive material has been delivered to a Florida customer.  While general licenses offer ease of acquisition, minimal documentation requirements and lower fees, they are subject to certain conditions, requirements, limitations and restrictions.  For additional information concerning issuance of a general license or the varying types of general licenses, review section 64E-5.204, F.A.C., and section 64E-5.206, F.A.C.

Radioactive materials whose nuclear properties, chemical and physical form, and activity present a greater health and safety hazard or are not otherwise permitted under a general license require issuance of a specific license.  The specific license is a document issued to an applicant, authorizing a particular use of radioactive material.  The license identifies the radioactive material, chemical and physical form, maximum activity and the purposes for which it may be used.  Applicants must demonstrate that they have appropriately trained and qualified personnel, appropriate facilities, equipment and procedures to ensure safe operations.  The fees for a specific license are significantly higher than a general license.  The specific license application process is detailed, requiring establishment and implementation of a detailed radiation protection program.

Florida allows authorization for a general license to be included under a specific license.  Applicants seeking to include authorization for generally licensed radioactive materials as part of their specific license must include a written request for general license authorization as part of their application; see section III.7. of this introduction for additional guidance.  Once approved, authorization for generally licensed radioactive materials will be included as part of Items 6, 7, 8 and 9 of the license, and a general license condition describing requirements associated with generally licensed radioactive materials will be included in the license.

This guide provides instructions for applicants seeking a specific license.  Unless otherwise noted, when used in this guide, the term “license” refers to a specific license.

Appendices, Exhibits and Supplements

Applicants must acquire equipment, train workers, and implement procedures that ensure compliance with regulatory requirements.  This guide provides a set of appendices, exhibits and supplements to assist in the development of a radiation protection program.

- Appendices are model procedures that may be used to address regulatory requirements.

- Exhibits are samples of the types of documents or forms that must be submitted as part of the application, and in some cases, are model forms that may be used to satisfy regulatory requirements.

- Supplements are documents used for preparation and submittal of an application.  Supplement A is Form DH-1054, which must be completed for each application.  Supplement B is a table of attachments that should also be completed and submitted with each application.  Supplement C is a checklist that should be used to ensure completeness of an application.

Model procedures and forms may be adopted by submitting them as part of the license application, or may be used as guides for developing equivalent procedures and forms.  Carefully review the regulations, model procedures and forms before deciding if the models are appropriate for the activities being requested.

IMPORTANT NOTICE:

The information provided in a license application must demonstrate that proposed equipment, facilities, personnel and procedures are adequate to protect public health and property in accordance with regulatory requirements.  Submission of incomplete or inadequate information will result in delays in the license approval process.  Additional information will be requested when necessary to ensure that an adequate radiation safety program has been established.  Such requests will delay completion of the application review, and may be minimized by a thorough study of the regulations and this guide prior to submitting the application.

While the adoption of the attached model procedures and forms should provide for a radiation protection program that complies with regulatory requirements, applicants may need to consider additional equipment, procedures and training that may be appropriate for the scope of their operations.

B.	APPLICABLE REGULATIONS

Florida is an Agreement State; it has an agreement with the U.S. Nuclear Regulatory Commission (NRC) to assume regulatory authority over most activities involving radioactive material within the state.  With certain exceptions, the Department of Health (department), Bureau of Radiation Control (BRC) regulates the possession and use of radioactive material within Florida.  Exceptions include nuclear power plants and federal agencies, and national security issues involving radioactive material, which remain under NRC jurisdiction.

Under authority of Chapter 404, Florida Statutes (the Florida Radiation Protection Act), the BRC issues licenses to users of radioactive material and performs inspections to ensure safe operations and compliance with Chapter 64E-5, Florida Administrative Code (F.A.C.), the department’s radiation control regulations.  Chapter 64E-5, F.A.C., is available on the Internet at http://www.doh.state.fl.us/environment/radiation.  The BRC amends these regulations periodically.  Licensees are notified of changes as they occur.  When applicable, licensees will need to revise their safety programs to address changes in regulatory requirements.

The parts of the regulations applicable to unsealed and/or sealed sources for use in in-vitro and clinical laboratory, academic, research and development, nuclear service and other miscellaneous uses subject to Part XIII are listed below, and should be used in conjunction with this guide.  Chapter 64E-5, F.A.C., Part XIII, Subparts A and D include rules that specifically apply to the use of unsealed and/or sealed sources.

Part I	“General Provisions”

Part II	“Licensing of Radioactive Materials”

Part III	“Standards for Protection Against Radiation”

Part IX	“Notices, Instructions and Reports to Workers; Inspections”

Part XIII	“Radiation Safety Requirements for Possession and Use of Sealed or Unsealed 
	 Sources of Radioactive Materials”

Part XV	“Transportation of Radioactive Materials”

Licensees engaging in transportation of radioactive materials or related activities are also subject to U.S. Department of Transportation (DOT) regulations, which are found in Title 49, Code of Federal Regulations (49 CFR), and are incorporated into Chapter 64E-5 by reference.  DOT regulations are available on the Internet at http://www.access.gpo.gov/nara/cfr/cfr-table-search.html#page1 and can be ordered from the U.S. Government Printing Office by calling (866) 512-1800 or writing P.O. Box 37954, Pittsburg, PA 15250-7954,  Attn: Superintendent of Documents.

C.	LICENSE REQUIREMENTS AND RESTRICTIONS

Licensees are required to confine use and possession of radioactive material to the locations and purposes authorized by the license.  The license is divided into two sections: Items and Conditions, which are described below.  The first section of the license lists Items 1 - 9.  The remainder of the document lists the license conditions, which may vary in number based on the authorizations provided by the license, but always begin with Condition 10.

License Items

| Item No. and Title | Description |
| --- | --- |
| 1. Name | Lists the legal name of the licensee (individual or business). If the license is issued to a business, Item 1 must list the company’s name as it is registered with the Florida Department of State, Division of Corporations; (800) 755-5111 or http://www.sunbiz.org/corpweb/inquiry/search.html. If a business operates under another name, Item 1 must list both the registered name and the fictitious name it is doing business as (d/b/a). |
| 2. Address | Lists the mailing address, which may be different from the physical address where records and radioactive materials are used/stored. If the two addresses are different, the physical address must be listed in License Condition 10; if they are the same, Condition 10 references the address listed in License Item 2. |
| 3. License Number | Lists the number assigned to the license by the BRC. The number should be referenced in all correspondence with the BRC. |
| 4. Expiration Date | Lists the date the license will expire. A radioactive materials license is valid for 5 years from the date of issuance. |
| 5. Category | Activities involving possession and use of radioactive materials are divided into license categories. Section 64E-5.204, F.A.C., lists license categories and fees. Conducting more than one category of licensed activity requires a separate license for each category of use. |

| 6. Radioactive Material | Describes the type (element and mass number) of radioactive material the license authorizes for possession and use. |
| --- | --- |
| 7. Form | Describes the form of radioactive material the license authorizes for possession and use (e.g., liquid, sealed source, manufacturer/distributor and model number). |
| 8. Possession Limit | Lists the maximum possession limit and if applicable, it identifies the quantity of sealed sources that the licensee may possess. Possession of quantities greater than are permitted by the license is a violation and may result in enforcement actions. |
| 9. Use | Describes the types of uses that are approved for the sources and devices listed in the previous items. Unauthorized use of radioactive material is a violation and may result in enforcement actions. |

License Conditions

License conditions describe requirements and limitations applicable to the radioactive materials authorized by the license.  Additional requirements and conditions may be incorporated as appropriate to protect public health and the environment.  If a licensee seeks added authorizations, supplementary license conditions may be added.  The following conditions may be applicable to unsealed and/or sealed sources for use in in-vitro and clinical laboratory, academic, research and development, nuclear service and other miscellaneous uses subject to Part XIII of Chapter 64E-5, F.A.C.

| - Authorized location of use and storage - Enforcement provisions - Authorized User (AU) and RSO designations - Radioactive material transfer limitations - Radioactive material transportation requirements - Enforcement provisions - Part III and IX provisions | - Bioassay requirements - Leak testing requirements - Inventory requirements - Contamination control program - Exemptions as applicable - Other restrictions or limitations - Licensee commitments |
| --- | --- |

II.	FILING AN APPLICATION

A.	GENERAL

Chapter 64E-5, F.A.C., this guide, forms, and other guidance documents are available on the BRC website: http://www.doh.state.fl.us/environment/radiation.

An application for a specific licenses authorizing unsealed and/or sealed sources for use in in-vitro and clinical laboratory, academic, research and development, nuclear service and other miscellaneous uses of radioactive material subject to Part XIII, must be submitted on Form 
DH-1054, "Application For Radioactive Materials License, Non-Human Use."  The form is included as Supplement A of this guide, and is also available on the BRC website.  Space provided on the application form is limited, so 8.5” x 11” paper should be used to append additional pages.  Each page submitted with the application should be identified and keyed to the item number on Form DH-1054 to which it applies.  Three copies of the application and all attachments must be submitted (original and two copies), with another copy retained by the applicant.  All application items must be addressed in sufficient detail to demonstrate that equipment, facilities, personnel qualifications and procedures are adequate to protect public health and safety or property.

Complete and submit the table provided as Supplement B to this guide to indicate whether model or equivalent procedures and forms have been included in the application.

| Mail to: | If using an overnight delivery service, use: |
| --- | --- |
| Department of Health Bureau of Radiation Control Radioactive Materials Program 4052 Bald Cypress Way, Bin C21 Tallahassee, FL 32399-1741 | Department of Health Bureau of Radiation Control Radioactive Materials Program 4042 Bald Cypress Way, Rm. 220.09 Tallahassee, FL 32399 |

With the exception of security-related information, all license applications and documents submitted to the BRC are available for review by the general public.  Do not submit proprietary information unless it is absolutely necessary for evaluation of the application.  Any request for withholding documents is subject to a determination by the department as to whether the document may actually be withheld in accordance with applicable laws and regulations.

Personal information about employees should not be submitted unless it is necessary.  Home addresses, home telephone numbers, dates of birth, and social security numbers should not be submitted unless the BRC specifically requests it.

When issued, the license will require that radioactive material be possessed and used in accordance with statements, representations and procedures provided in the application and supporting documentation (which are incorporated by referenced into the license).  Regulatory requirements specified in Chapter 64E-5, F.A.C., shall govern unless the statements, representations and procedures set forth in the license application and correspondence are more restrictive than the regulations.

B.	LICENSE FEES

The following fees are assessed:

Application fee	A non-refundable fee for processing the license application.  The amount is dependent on the category of license the applicant is seeking; refer to section 64E-5.204, F.A.C., or Regulatory Guide 6.20 for a description of application fees.  Review of the application will not begin until the proper fee is received by the department.  An application fee is also required to process an application for a new license replacing an existing license due to a change of ownership.

Annual fee	An annual fee covers department costs for administration of the materials licensing program.  The amount is dependent on the license category.  Refer to section 64E-5.204, F.A.C., or Regulatory Guide 6.20 for a description of annual fees.  Annual fees are due within 60 days of issuance of the new license; an invoice for this fee is included with the cover letter accompanying a new license.

Reclamation fee	In addition to the application and annual fees, a reclamation fee will be assessed for the Radiation Protection Trust Fund, established to pay department costs associated with a licensee’s abandonment of radioactive materials, default on lawful obligations, insolvency, or other inability to meet regulatory requirements, and to assure the protection of the public and environment.  Reclamation fees are equal to 5% of the annual fee.  Reclamation fees are due within 60 days of issuance of a new license; a fee invoice is included with the cover letter accompanying a new license.

Notes:	1.	Annual and reclamation fees are assessed on the anniversary of the license issuance date.  An invoice is sent to the licensee 30 days in advance of the due date.

2.	Fees are not assessed for license renewals, amendment requests, licensing actions, inspections initiated by the department, license terminations, or requests for regulatory information (except for document copying costs).

III.	CONTENTS OF AN APPLICATION

This section provides instructions on completing each item listed in Form DH-1054.

1.a.	NAME AND MAILING ADDRESS

List the name and mailing address of the individual or company to whom the license will be issued.  An applicant corporation or other legal entity must be listed by the legal name registered with the Department of State, Division of Corporations.  If a fictitious name is to be included, it will be identified as the name the applicant is doing business as (d/b/a) and must also be registered with the Division of Corporations.  For example, “ABC Corporation d/b/a ABC Enterprises of Florida.”  Business registration may be verified by contacting the Division of Corporations at (800) 755-5111 or on the Internet at: http://www.sunbiz.org/corpweb/inquiry/search.html.

To assist in identifying the proper legal entity, applicants should also list their business’ Federal Employer Identification (FEI) or Document Number if known or applicable; the FEI or Document Number is available on the Division of Corporations website.

If the mailing address is different from where radioactive material will be used and/or stored and where license-related records will be maintained, list that address in Item 1.b.

1.b.	STREET ADDRESS AT WHICH RADIOACTIVE MATERIAL WILL BE USED and/or stored, IF DIFFERENT FROM 1.a.

Identify by street address the location where radioactive material will be used and/or stored and where license-related records will be maintained, if different that the address listed in Item 1.a.  Non-contiguous locations of use and storage may require a separate license; refer to section 
64E-5.213, F.A.C.

2.a.	LICENSE FEE CATEGORY

Indicate the appropriate license fee category, as follows:  For in-vitro and clinical laboratory, list category 3L(II); for academic, list category 3L(III); for industrial or medical research and development, list category 3K; for nuclear service, list category 3P; for gas chromatography devices, list category 3N; for reference or calibration sources equal to or less than one millicurie total, list category 3O; for all other specific license except otherwise noted, list category 3L(V).  Refer to section 64E-5.204, F.A.C., or to Regulatory Guide 6.20 for a listing of fees.

2.b.	LICENSE FEE ENCLOSED

Indicate the amount of the enclosed license application fee in the space that is provided.  Refer to section II.B., of this guide, for a description of licensing fees.

3.	THIS IS AN APPLICATION FOR:

Mark the appropriate choice; if submitting an amendment request or a renewal application, indicate the applicable radioactive materials license number in the space provided.

4.	INDIVIDUAL USERS

List each individual to be designated as an authorized user (AU) of radioactive material.  Specify the radioactive material that each individual is qualified to use, (i.e., sealed and/or unsealed sources) and describe the extent to which each individual is authorized to use each source type or device.  Condition 12 of the license will either list each individual with their corresponding authorizations or it may state that “licensed materials shall be used by, or under the supervision and in the physical presence of, individuals who have successfully completed the licensee’s training program for authorized users described in their license application. . . “  Maintaining documentation of training is required to demonstrate that personnel are adequately trained.

The training program (submitted as part of the radiation protection program under III.13.) must provide a commitment that all AUs will complete either:

- An approved radiation safety course provided by a third party (manufacturer/distributor or another approved training provider), that may be supplemented with supervised  experience using the licensed radioactive material and instruction in the licensee’s operating and emergency (O&E) procedures; or

- An approved in-house training program meeting the requirement of section 64E-5.1307, F.A.C.  If this option is chosen, a detailed description of the in-house training program must be submitted.

5.	RADIATION SAFETY OFFICER (RSO)

Provide the name of the individual assigned the position of RSO.  This person is designated by and responsible to management for implementation of the radiation safety program and for ensuring compliance with the applicable regulations and license provisions.  As a minimum, the RSO must have sufficient training and experience to be an authorized user of the requested radioactive materials.  Additional training in administration of a radiation protection program is recommended for the RSO position.

6.	TRAINING AND EXPERIENCE IN RADIATION SAFETY

a.	FORMAL TRAINING IN RADIATION SAFETY

Submit documentation of radiation safety training for each individual listed in Items 4 and 5 of the application.  Restrict training documentation to relevant information (i.e., Information demonstrating that the individual has the radiation safety training and experience specific to the requested activities to be conducted).  Appropriate training certificates such as those provided by the manufacturer/distributor or other approved third parties are acceptable.  Such certificates may need to be supplemented with documentation of training in the applicant’s operating and emergency (O&E) procedures to satisfy Chapter 64E-5, F.A.C., Part XIII training requirements, because third party trainers may not provide such training.

b.	EXPERIENCE

Describe any additional relevant work experience with radiation and where and under whose supervision the experience was obtained.  Identify the radioactive material, chemical and physical form, maximum activity, and the tests, procedures or research conducted using the material.  A curriculum vitae may also be included as supporting documentation, but please do not include the individuals’ birth dates or social security numbers.

7.	RADIOACTIVE MATERIAL

a.	ELEMENT AND MASS NUMBER

List each type of radioactive material requested; refer to the example provided below.

b.	CHEMICAL AND/OR PHYSICAL FORM

Complete for each type of radioactive material requested.  Identify if the radioactive material, is a solid, liquid, aerosol or gas.  Or you may request authorization for any form, or any form except gas.  Describe the chemical form (e.g., inorganic or organic).  For sealed sources, state the name of the source, manufacturer and the source model number; refer to the example provided below.

c.	MAXIMUM AMOUNT TO BE POSSESSED AT ANY ONE TIME

Complete for each radioactive material requested.  List the maximum activity of radioactive material that may be possessed.  For sealed sources, indicate the total number of sources and maximum activity per source.

Example:

| (a) ELEMENT AND MASS NUMBER | (b) CHEMICAL AND/OR PHYSICAL FORM | (c) MAXIMUM AMOUNT TO BE POSSESSED AT ANY ONE TIME |
| --- | --- | --- |
| 1. Cesium 137 | Sealed source (A.B.C. and Associates Model Model 58-2) | 1 source; not to exceed 330 millicuries |
| 2. Iodine 125 | Liquid humic acid in a phosphate buffered solution | 200 millicuries |
| 3. Thorium 230 | Any form, except gas | 0.1 microcurie |
| 4. Nickel 63 | Foil source (Alpha Beta Chi Corporation Model XYZ-123 or Delta Kappa Rho Company Model 987) | 2 sources; not to exceed 15 millicuries each |
| 5. Sulfur 35 | Any form, except gas | 10 millicuries |

If authorization for generally licensed (GL) sources or devices is sought, include a request for GL sources and devices.

Example:

| (a) ELEMENT AND MASS NUMBER | (b) CHEMICAL AND/OR PHYSICAL FORM | (c) MAXIMUM AMOUNT TO BE POSSESSED AT ANY ONE TIME |
| --- | --- | --- |
| 6. Radioactive material Sealed or contained source distributed to a general licensee per 64E-5.206(1) & (4), F.A.C. |  | No single source to exceed that quantity authorized for the general license |

Note:	Applicants seeking authorization for GL sources and devices must comply with section 64E-5.1308, F.A.C., which describes requirements for GL devices possessed under a specific license.  Annual inspections/inventories must include the GL sources, and the sources must be leak tested at the interval specified by the manufacturer.

8.	PURPOSE FOR WHICH RADIOACTIVE MATERIALS WILL BE USED

Complete for each radioactive material requested.  Include the name of the manufacturer and model of the device or source holder in which each source is used or stored.

Example:

1.	To be used in a low dose irradiator for the calibration of radiation detection instruments.

2.	To be used for the measurement of the uptake of humic acids by phytoplankton.

3.	To be used as reference, calibration and/or sample spike sources for analysis of environmental samples.

4.	To be used in Psi-Epsilon Model 900-0822 electron capture detector (ECD) cells in Psi-Epsilon  Model 1230C gas chromatographs (GCs) for sample analysis.

5.	To be used for in vitro testing.

If authorization for GL material, sources, or devices is requested, describe the intended use of the GL material, sources or devices.

Example:

6.	To be used in devices approved for receipt under general license provisions.

CURRENT INVENTORY

Applicants for renewal of an existing license must include an inventory of all sealed sources and devices currently possessed.  List all generally licensed and specifically licensed sources and devices, and indicate the licensing designations for each (i.e., general or specific).  If in possession of any exempt sources (e.g., check sources), it is recommended that they be included in the inventory in order to avoid any confusion over their licensing status.  If exempt sources are included, identify their exempt status on the inventory form.

9.	RADIATION DETECTION INSTRUMENTS

List and describe all radiation detection instruments used for contamination control and other required surveys, analytical analysis or research, including those used for analysis of leak tests and iv-vivo or in-vitro bioassays.  An instrument shall be available for detection of radioactive materials authorized by the licensee.  Provide all requested information for each instrument as specified and as required by section 64E-5.1318, F.A.C.  Attachments may be included as necessary.

10.	CALIBRATION OF INSTRUMENTS

If radiation detection instruments will be used, mark the appropriate box to indicate who will perform calibrations.  Indicate the calibration frequency for all radiation detection instruments.  Note that radiation detection instruments and equipment used for quantitative measurements must be calibrated at least annually.

a.	CALIBRATED BY SERVICE COMPANY

If a service company will be used, list the vendor(s) name, address, license number and the government agency that issued the company’s license (i.e., NRC or a state agency).  Survey instruments must be calibrated at least annually per subsection 64E-5.314(2), F.A.C.

b.	CALIBRATED BY APPLICANT

If seeking approval to calibrate instruments in-house, submit detailed information describing the facilities, equipment, personnel, and procedures to be used to perform the calibrations.  Contact the BRC for additional guidance on equipment calibration requirements.  Note: in-house calibration requires use of reference sources; list each requested calibration/reference source in Item 7.

11.	PERSONNEL MONITORING DEVICES

Common personnel monitoring badges include film badges, thermoluminescent dosimeters (TLDs) and optically stimulated luminescent dosimeters (OSLDs), which are described below.  Personnel monitoring badges must be capable of detecting the type of radiation (e.g., beta, gamma, neutron) emitted by the radioactive material authorized by the license.

Badge processors must hold accreditation form the National Voluntary Laboratory Accreditation Program (NVLAP) of the National Institute of Standards and Technology.  A list of NVLAP-accredited badge vendors is available at http://ts.nist.gov/Standards/scopes/dosim.htm.  A list of commercial personnel monitoring badge vendors is available from the BRC upon request.

Each order of badges includes a control badge for measuring the amount of background radiation the badges receive each monitoring period.  This enables background radiation to be subtracted from the total reading to provide an accurate record of each worker’s occupational exposure.  When not in use, personnel monitoring badges should be stored with the control badge in an environment protected from radiation, chemicals, excessive heat, light, moisture, etc. to ensure accurate dose records.  The control badge must be returned with the other personnel monitoring badges each monitoring period.

Film badges are small pieces of x-ray film contained in a plastic holder.  The film darkens in proportion to the amount of radiation it has been exposed to, so the film’s optical density provides a measurement of the wearer’s radiation exposure.  Film badges must be exchanged monthly.

TLDs are personnel monitoring badges that contain small crystals capable of storing some of the energy from radiation.  Heating the crystals releases the stored energy as light.  The amount of light released is proportional to the amount of radiation the TLD badge received, which is measured to determine the badge wearer’s dose.  TLDs must be exchanged at least every three months.

OSLDs measure radiation using a thin layer of aluminum oxide.  A laser light stimulates the aluminum oxide after use, causing it to become luminescent in proportion to the amount of radiation exposure.  OSLDs must be exchanged at least every three months.

in accordance with section 64E-5.315, F.A.C., and subsection 64E-5.1310(4), F.A.C., a whole body film badge, optically stimulated luminescent dosimeter (OSLD), or thermo luminescent dosimeter (TLD) personnel monitoring device shall be supplied and required to be worn by:

- Any individual likely to receive a radiation dose exceeding 500 millirem in any one year; or

- Any individual using or assisting in the use of unsealed sources of radioactive materials of any gamma-emitting isotope with a gamma ray energy greater than 50 kiloelectron volts or the use of any beta-emitting isotope with a maximum beta energy of 300 kiloelectron volts or more.

In accordance with subsection 64E-5.1310(5), F.A.C., an extremity film badge, OSLD, or TLD, monitoring device shall be worn by:

- Any individual using or assisting in the use of unsealed sources of radioactive materials of 1,000 microcuries or more of beta-emitting isotopes with a maximum beta energy of 1,000 kiloelectron volts or more in any month; or

- Any individual using or assisting in the use of unsealed sources who receives a dose of 40 millirem or more on a whole body film badge, OSLD, or TLD, for 2 consecutive months.

12.	FACILITIES AND EQUIPMENT

Facilities and equipment must be adequate to protect health and minimize danger to life and property.

Facility Diagram

Submit an annotated diagram of the facility identifying all areas where radioactive materials are received, stored (including waste), prepared, measured and used.  Depict the countertops, glove boxes, fume hoods, sinks, cabinets, refrigerators, freezers and other areas where radioactive materials are used, disposed or stored.  Identify all adjacent areas and occupied workstations.

Security

Radioactive materials must be secured to prevent unauthorized access or removal.  Describe the security measures in place to prevent unauthorized access or removal of radioactive materials.

Fume Hoods and Glove Boxes

Specify the withdrawal rate for fume hoods (an average linear face velocity of 150 feet per minute with a minimum of 125 feet per minute is recommended) and indicate the appropriate window sash height required to maintain this withdrawal rate.  If more than one fume hood is available, depict their proximity to each other and indicate if they are used in tandem.  A charcoal or high efficiency particulate air (HEPA) filter may be installed in the fume hood or glove box to maintain effluent releases within allowable limits.  If a filtration device is used, indicate the frequency at which is monitored and exchanged.  Submit applicable procedures.

Ventilation Diagrams

If using radioactive gases outside of a properly functioning fume hood or glove box, submit additional details relating to room ventilation indicating the locations of each supply, return and exhaust vents and the withdrawal rate in cubic feet/minute.  As applicable, submit calculations demonstrating compliance with the permissible annual limit on intake, derived air concentrations and effluent concentrations.

Safety Equipment

Submit a list of available safety equipment (e.g., remote handling equipment, personnel respiratory protection equipment, shielded containers, spill kits and fixed area monitors), available shielding and describe the methods used for minimizing the spread of contamination from unsealed sources (e.g., installing vinyl flooring in lieu of carpet and using plastic backed absorbent paper on countertops).

Instrument Calibration Shielding, Equipment and Configuration

If performing calibrations of portable radiation detection instruments, describe the related shielding, equipment (e.g., shielded box, irradiator chamber, remote viewing and handling equipment) and facilities utilized.

13.	RADIATION PROTECTION PROGRAM

Submit a detailed description of the proposed radiation protection program, which must include the following components.  The appendices and exhibits included with this guide are model procedures and forms that may be adopted by including them as part of the submitted radiation protection program, or used as guides for developing equivalent procedures and forms.

Reminder: Complete Tables 1 – 3 of Supplement B of this guide to indicate whether model or equivalent procedures and forms have been submitted, and attach a copy to Form DH-1054.

Member of the Public (MOP) Dose Limit Compliance Study	Appendix A

Applicants for a new license must submit proposed procedures for demonstrating compliance with the public dose limits specified in section 64E-5.313, F.A.C. (2 millirem in any one hour and 100 millirem per year).  Applicants for renewal of an existing license must submit a completed study demonstrating compliance with the limits.  Appendix A is a model study.

ALARA Policy	Appendix B

Applicants must submit a policy describing management's commitment to the ALARA philosophy of maintaining doses as low as reasonably achievable, and a description of the commitments of management and workers for implementing the policy.  Appendix B is a model ALARA policy.

Duties and Responsibilities of the Radiation Safety Officer (RSO)	Appendix C

Applicants must submit a description of the RSO’s duties and responsibilities that includes the duties listed in section 64E-5.1305, F.A.C.  Appendix C is a model procedure.

Radiation Safety Training Program	Appendix D

Multiple training requirements are applicable to unsealed and/or sealed sources used in-vitro and clinical laboratory, academic, research and development, nuclear service and other miscellaneous uses subject to Part XIII of Chapter 64E-5, F.A.C.

- Radiation awareness training (“Instructions to Workers”) must be provided to personnel engaged in licensed activities (AUs and workers under their supervision).  This training is also recommended for individuals that work in the vicinity of radioactive materials.  Section 64E-5.902, F.A.C., specifies radiation awareness training requirements.  The rule does not specify the minimum duration for this training, because the amount of training needed will vary depending on the scope of the radiological hazards present in the applicant’s workplace.

- Hazmat employee training must be provided to any worker with job functions associated with radioactive materials.  Hazmat employee training is specified in 49 CFR Part 172, Subpart H.  U.S. DOT regulations are incorporated by reference in Chapter 64E-5, F.A.C.  The rule does not specify the minimum duration for this training.

- Authorized User (AU) training must be provided to workers independently working with radioactive materials or supervising such activities by other workers.  AU training requirements are specified in subsection 64E-5.1307, F.A.C.

The training program must describe how the above training requirements will be addressed.  Because the topics that must be addressed to satisfy radiation awareness and hazmat employee training overlap with the topics that must be covered during AU training, these training requirements may be addressed concurrently.  Training can be provided by qualified third parties, in-house, or by using a combination of the two.  Appendix D is a model radiation safety training program aimed at licensees using an approved third party training course, supplemented by in-house training in the licensee’s O&E procedures.

Applicants seeking to conduct in-house training must address the requirements specified in subsection 64E-5.1307(2), F.A.C., by submitting a detailed description of the training program for review and approval, including a description of training reference materials, instructor qualifications, and a sample exam.

Personnel Monitoring Procedure	Appendix E

If PM is conducted, a procedure is required to provide instructions on proper use, exchanges, use of spare badges, lost or damaged badges, and PM record-keeping requirements.  Appendix E is a model PM procedure.  Exhibits E, F and G are model forms for addressing declared pregnant female requirements.

Operating and Emergency Procedures

Sections 64E-5.208 and 64E-5.1302, F.A.C., require establishment and implementation of O&E procedures that provide instructions adequate to ensure safety to workers, the public and to property.  As a minimum, O&E procedures must include the procedures described below.

- Operating Procedure	Appendix F

Due to the varying nature and scope of licensed operations for which this regulatory guide was developed, it is not feasible to provide a model operating procedure that addresses all uses of unsealed and/or sealed sources in in-vitro and clinical laboratory, academic, research and development, nuclear service and other miscellaneous uses subject to Part XIII of Chapter 
64E-5, F.A.C.  Applicants must consider their own unique licensed operations and use of unsealed and/or sealed sources to ensure that all necessary instructions are included.  Therefore, we suggest that you supplement our model procedure with other applicable procedures that relate to your own unique operations.

Appendix F provides generic instructions on availability of procedures, general rules of use/
ALARA principles, radiation surveys, security and routine maintenance.  Operating procedures must include instructions on those topics; additional guidance is provided below.

Training

Proper training is the most important factor contributing to safe handling, transport and use of radioactive materials.  Individuals shall complete applicable training prior to handling or using radioactive materials.

Availability of Procedures

Operating procedures must include a commitment that AUs will have access to the licensee’s O&E procedures, as well as the manufacturer operation/maintenance manual for each type of licensed device or instrument possessed.

Personnel Monitoring Instructions to Workers

Individuals supplied whole body and/or extremity personnel monitoring devices must be provided instructions about when and how the device should be worn, stored and exchanged.

General Rules of Use/ALARA Principles

Instructions should be provided on techniques for minimizing dose, general precautions and the performance of radiation surveys if damage to a radioactive source or device is suspected.  The instructions may be supplemented by any specific instructions provided by the manufacturer or distributor of the radioactive source or device.

Security

Radioactive materials must be stored and used in a manner that secures them from unauthorized access or removal.  Additional controls (e.g., building locks, monitored security systems, fences, guards, etc.) should be utilized as appropriate to enhance security.

Routine Maintenance

Radioactive sources and devices may require periodic maintenance.  Maintenance must be performed in accordance the manufacturer’s procedures and recommendations.  Typically, the manufacturer’s procedures allow for the performance of routine maintenance by authorized users.  The performance of non-routine maintenance may be restricted to those acting as the manufacturer’s representative.

Posting Requirements

Areas where radioactive materials are used and stored must be posted with appropriate radiation warning as described in 64E-5.323, F.A.C.  This procedure also addresses the posting of required documents as specified in section 64E-5.901, F.A.C.

- Procedures for Ordering, Receiving, Opening and Shipping	Appendix G
  Packages Containing Radioactive Material

Procedures must address preparation and handling of incoming and outgoing shipments of radioactive material transported by common carriers, and if applicable, for radioactive materials transported as private use shipments.  The instructions must conform to current U.S. DOT regulations.  Sample shipping papers and emergency response information must be provided.  Appendix G is a model procedure, Exhibit C is a sample shipping paper for private use shipments, and Exhibit D is a model emergency response information sheet.

- Rules of Use of Unsealed Sources	Appendix H

If using unsealed radioactive materials, the procedures must provide instructions to personnel addressing the safe use and handling of unsealed sources of radioactive material.  Appendix G is a model laboratory rules of use procedure.

- Contamination Control	Appendix I

If using unsealed radioactive materials, the procedures must provide instructions to personnel addressing the control of the spread of contamination.  Appendix H is a model contamination control procedure.

- In-Vivo Thyroid Bioassay	Appendix J

If using unsealed radioiodine, the procedures must address in-vivo thyroid bioassay monitoring to assess the internal exposure resulting from the use of unsealed radioisotopes.  Table 1 of the U.S. Nuclear Regulatory Commission (NRC) Regulatory Guide 8.20, “Applications of Bioassay for I-125 and I-131,” specifies the criteria for the establishment of a bioassay monitoring program.

- In-Vitro Urinalysis Bioassay	Appendix K

If using hyrdorgen-3 (H-3) in tritiated compounds, including HTO, the procedures must address in-vitro urinalysis bioassay monitoring to assess the internal exposure resulting from the use of unsealed radioisotopes.  Establishment of a bioassay monitoring program must be in accordance with Table 1 of the U.S. Nuclear Regulatory Commission (NRC) Regulatory Guide 8.32, “Criteria for Establishing a Tritium Bioassay Program.”

- Emergency Procedures	Appendix L

The procedures must provide instructions for responding to accidents, unusual events and the loss or theft of radioactive material and must include emergency notification numbers for the RSO and BRC.  If using unsealed sources, the procedures must also address the incidental release or spills of radioactive aerosols, gases, fine particulates, liquids and solids and must include emergency notification numbers for the RSO and BRC.  Appendix J is a model emergency procedure.

Leak Testing Procedures

Sealed sources must be tested at regular intervals to ensure that the radioactive material is not leaking contamination.  Leak test requirements are specified in section 64E-5.1303, F.A.C.

- Leak Test Frequency, Collection and Records	Appendix M

Sealed sources must be tested at regular intervals to ensure that the radioactive material is not leaking contamination.  Leak test requirements are specified in section 64E-5.1303, F.A.C.  If a leak test kit is used, indicate the manufacturer and model number.  The procedure must specify the leak test interval and include instructions on collecting leak test samples.  This information may be obtained from the manufacturer.  The procedures must identify the vendors who will be contracted to conduct leak test analysis or include a commitment to have analysis performed by licensed vendors.  The procedures must provide instructions for retention of leak test records.  Appendix M is a model leak test procedure.

- Leak Test Analysis	Appendix N

If authorization to perform in-house leak test analysis is sought, additional procedures addressing sample analysis must be submitted.  The procedure must describe the instrumentation used to perform the analysis and include step-by-step instructions for calculating instrument efficiency and performing sample counts.  Appendix N is a model leak test analysis procedure.

Inventory Procedure	Appendix O

The procedure must provide instructions for performing annual physical inventories and inspections of all generally and specifically licensed sealed source or devices.  Inspections must evaluate the physical condition of the sealed source or device and the associated labels.  
A sample inventory form is also required.  Appendix O is a model inventory procedure; Exhibit A is a model inventory form.

Record Retention Procedure	Appendix P

Certain records must be retained for specified periods of time for compliance purposes.  These intervals have been established for BRC inspection staff and other authorized entities (e.g., U.S. DOT) to have access to the documents as required by the regulations.  Appendix N provides a model procedure addressing record retention requirements.

Notification and Reporting Procedure	Appendix Q

Notification and reporting requirements are specified in Parts II and III of Chapter 64E-5, F.A.C.  Appendix Q provides a model procedure summarizing notification and reporting requirements.

Radiation Detection Instrumentation Calibration Procedure	Appendix R

Radiation detection instruments used for quantitative analysis (e.g., dose rate and effluent monitoring) must be calibrated annually for each type of radiation measured.  Sections 64E-5.314 and 1318, Florida Administrative Code (F.A.C.), provide additional information regarding radiation detection instrument specifications and calibration.  Appendix R is a model radiation detection instrument calibration procedure.

14.	WASTE DISPOSAL	Appendix S

Submit a procedure describing how radioactive materials will be disposed.  The procedure must include a commitment that sealed sources will be disposed of either by returning it to the manufacturer or by transferring it to a specifically licensed recipient.  Please note that low-level radioactive waste brokers and most manufacturers require a fee to accept radioactive materials for disposal.  While certain limits and restrictions apply, unsealed sources may also be disposed by effluent and evaporative release and by decay-in-storage.  Appendix S is a model waste transfer/disposal procedure.

15.	CERTIFICATE

A radioactive materials license is a legal document.  License applications and license-related correspondence must be signed and dated by an individual (certifying official) authorized to make legally binding statements for the applicant.  Examples of positions that are recognized as certifying officials include owner, president, vice president, chief executive officer, chief operating officer, etc.  Positions that are not recognized as certifying officials include RSO, environmental health & safety director, and laboratory manager.

A certifying official may delegate authority to make legally binding statements to specific individuals or positions (e.g., manager, director, RSO) by submitting a written statement authorizing the delegation.  Exhibit H is a model form for documenting a delegation of authority.

IV.	LICENSE AMENDMENTS

Licensees are required to conduct operations in accordance with applicable regulations and the statements, representations and procedures contained in the license application and supporting documents.  The license must be amended if any changes are planned.  Submittal of an amendment request does not allow immediate implementation of proposed changes.  Until the license has been amended to reflect approval of the change(s), the licensee must comply with the original terms and conditions of the license.

Applications for license amendments may be submitted in letter form or on Form DH-1054, “Application For Radioactive Materials License, Non-Human Use.”  The request must be dated and signed by a certifying official, identify the license by name and number, be submitted in triplicate, and clearly describe the nature of the changes, additions or deletions requested.  References to previously submitted documents must be specific and identify the applicable information by date, page and paragraph.  The licensee must maintain a copy of the submitted and referenced documentation on file for inspection.  To prevent the potential for identity theft, do not submit documentation that lists individuals’ social security numbers or birth dates.

V.	LICENSE RENEWAL

Absent any actions by the department or the licensee, a license remains in effect for five years.  An application for license renewal must be received by the department at least 30 days prior to the expiration date.  This filing will ensure that the license does not expire until final action on the application has been taken, as provided for by subsection 64E-5.207(3), F.A.C.  If the application is received less than 30 days before the expiration date, the facility or individual may be without a valid license when the license expires.  Renewal applications should be filed using Form DH-1054, “Application For Radioactive Materials License, Non-Human Use” (Supplement A of this guide).  Renewals require submittal of an entire new application, completed as if it were an application for a new license, with complete and up-to-date information about the applicant's radiation protection program, demonstrating compliance with all licensing and regulatory requirements in effect at the time of renewal.  Renewal applications should be submitted without reference to documentation and information submitted previously.  To prevent the potential for identity theft, never submit documentation that lists individuals’ social security numbers or birth dates.

VI.	LICENSE TERMINATION

Prior to license termination, the licensee must properly dispose of all licensed radioactive material possessed.  Complete Form DH-1059, “Certificate – Disposition of Radioactive Material” to satisfy the requirements of section 64E-5.214, F.A.C., and submit it to the Bureau before the expiration date of the license with a request that the license be terminated.

1.50 Intro  7-19-13