I.	Introduction

Section 64E-5.312, Florida Administrative Code (F.A.C.), requires licensed operations to be conducted so that the following limits are met:

- Radiation doses in unrestricted areas do not exceed 2 millirem (0.02 mSv) in any one hour

- Doses to members of the public do not exceed 100 millirem (1 mSv) in a year

Section 64E-5.313, F.A.C., requires surveys, calculations and/or environmental monitoring to be used to demonstrate compliance with the dose limits.  A member of the public (MOP) dose compliance study (MOP study) provides documentation of compliance with both regulatory limits.  This procedure describes methodologies developed by the Florida Bureau of Radiation Control (BRC), intended for use by licensees authorized to possess and use unsealed and special form sources of radiation, to demonstrate compliance with the MOP dose limits.

The below marked box indicates how this procedure has been utilized:

|  | New license applicant: the procedure describes the methodology that will be used to conduct the MOP study after licensed activities begin. |
| --- | --- |
|  | Renewal application: the procedure describes the methodology and results of the completed MOP study of existing operations. |

II.	Dose Limit for Unrestricted Areas

For typical academic, in-vitro, clinical and environmental laboratories, academic, research and development, nuclear service, and other miscellaneous uses subject to Part XIII of Chapter 64E-5, F.A.C., there are generally three situations that must be addressed to demonstrate compliance with the 2 millirem in any one hour dose limit for unrestricted areas.

- Use and storage of licensed materials at approved locations;

- Transfer and storage of radioactive waste; and

- Storage of licensed materials in transport vehicles, if applicable;

Section A demonstrates compliance with the unrestricted area dose limit for shipment of licensed materials to and from approved locations of use and storage.

Section B’s Method 1 describes the procedure followed when a survey meter is available to conduct radiation measurements.  Compliance with the unrestricted area dose limit can also be demonstrated without direct measurements.  Section B’s Method 2 describes the procedure followed when a survey meter is unavailable.  The marked box indicates the method selected for use in this study.

A.	Transport Vehicles and Temporary Job Sites

Security procedures approved by the BRC and incorporated into the license describe measures taken by authorized users, and approved personnel working under their supervision, to ensure that public access to restricted areas, incident to the transport of licensed materials to approved locations of use and storage is prevented.  During transport and storage at approved locations, procedures require licensed materials to be secured from public access or removal.  While in use, licensed materials are maintained under the direct supervision of the authorized user to prevent unauthorized access.  Adherence to these procedures ensures compliance with the 2 mrem in any one hour public dose limit.

II.	Dose Limit for Unrestricted Areas

B.	Permanent Facility

|  | Method 1. Physical Surveys |
| --- | --- |

Procedures approved by the BRC prevent unauthorized public access to licensed materials at the permanent facility.  When stored at other approved locations of use, as applicable, licensed materials are secured in an approved storage area.  When in use, licensed materials will be continuously observed by the authorized user, or approved personnel working under his/her supervision.

A radiation detection instrument was used to measure ambient radiation levels in the unrestricted areas around the permanent and waste storage area while all possessed licensed radioactive material were in storage.  This survey evaluated the “worst case scenario” – when radiation emitted from licensed operations was at their highest levels.  Survey results revealing dose rates below 2 millirem per hour demonstrate compliance.

The following information is attached:

- Date of the survey and the name of the individual(s) performing the measurements.

- Information about the instrument used to perform the survey (manufacturer and model number, the types of radiation detected by the instrument, its minimum and maximum range, and the date it was last calibrated).

- Diagram of the permanent facility identifying the location of use and storage, including waste, adjacent unrestricted areas, nearby MOP workstations, and the locations where all recorded measurements were taken.

- Information about the type and quantity of licensed radioactive materials present during the survey and a description of their location and orientation within the storage area.

- Results of survey(s) of unrestricted area radiation levels, with results keyed to facility diagram.

Note:	If surveys note radiation levels > 2 mR/hr, attach a description of controls in place to further restrict access to the storage area (e.g., establishment of expanded restricted area around the storage area, using barricades and/or posted notices).

|  | Method 2. Calculations |
| --- | --- |

Radiation levels in unrestricted areas can be calculated using information provided by the respective manufacturers and/or distributors.  Prior to shipment, the manufacturer and/or distributor lists the Transport Index (TI) number on the RADIOACTIVE YELLOW II labels on the transport container’s exterior surface.  The TI indicates the radiation levels at 1 meter (3.3 feet) from the container when it contains the radiation source.  The TI value was used as the basis for the calculations.  If the TI is less than 2, then radiation levels in all directions around the radiation source when it is stored in its transport container are 2 millirem per hour (or less) at 1 meter, so that is the boundary of the restricted area.  In the same way, it may be presumed that the dose rate of transport containers bearing a WHITE I label, have a dose rate less than 1 millirem per hour at 1 meter.   Additional distance and shielding provided by the storage area lower the dose rate even further.  Storing the source in its shipping container, then storing the container in a cabinet, locker, room, etc. prevents unauthorized access to within a meter or more, so no MOP can receive 2 millirem in any one hour.  When calculating for two or more radiation sources,

II.	Dose Limit for Unrestricted Areas

B.	Permanent Facility Method 2: Calculations    (Continued)

each source container’s TI is added together.  While this method is overly conservative, it serves to ensure that the 2 mrem limit is not exceeded.

The following information is attached:

- Diagram of the permanent facility identifying the restricted area, adjacent unrestricted areas, and distance to MOP workstations.

- Information about the type and quantity of radioactive materials present and a description of the location and orientation within the storage room.

- Copies of manufacturer-provided documentation providing information on dose rates and/or TI numbers for the radiation sources being evaluated.

- Results of calculations demonstrating estimated radiation levels in unrestricted areas, with results keyed to the facility diagram.

Note:	If calculations note radiation levels > 2 mR/hr, attach a description of controls in place to further restrict access to the storage area (e.g., establishment of expanded restricted area around the storage area, using barricades and/or posted notices).

III.	Annual Public Dose Limit

A.	Total Effective Dose Equivalent

Total Effective Dose Equivalent (TEDE) describes the dose from summation of internal and external radiation doses.  The internal radiation dose, or Committed Effective Dose Equivalent (CEDE), and the external radiation dose, Deep Dose Equivalent (DDE), are determined separately using the following supplemental sheets.

- If licensed only for sealed sources of radioactive material (RAM), complete Table 1.

- If licensed for unsealed RAM, or both sealed sources and unsealed RAM, complete Table 2, using Method 1 or 2.

- Include all additional attachments as appropriate.

- Sum the millirem values for the TEDE boxes as directed.  If a RAM use is not among the listed categories, an independent assessment of its contribution to the MOP dose must be made.  Attach documentation of all dose calculations performed to complete the assessments.

- *RAM effluent values are listed in “State of Florida Bureau of Radiation Control ALIs, DACs, and Effluent Concentrations,” July 1993; this document is appended to the end of Chapter 64E-5, F.A.C.

Definitions and Acronyms referenced above and the attached supplemental sheets are defined in section 64E-5.101, F.A.C.

III.	Annual Public Dose Limit

A.	Total Effective Dose Equivalent     (Continued)

| Table 1. SEALED SOURCES |  |  |
| --- | --- | --- |
|  | TEDE (mrem) | - Complete Section III.B. [Deep Dose Equivalent (DDE)] - Enter the calculated DDE value in the space provided to the left; for external  doses from sealed sources; DDE = TEDE Note: To demonstrate compliance with the yearly MOP dose limit specified in 64E-5.312(1)(a), the TEDE must be < 100 mrem. |

| Table 2. UNSEALED RAM or SEALED SOURCES & UNSEALED RAM |  |  |
| --- | --- | --- |
|  |  | Method 1. |
|  | DDE (mrem) | - Complete Section III.B. [Deep Dose Equivalent (DDE)]  - Enter the calculated DDE in the space provided to the left |
|  | CEDE (mrem) | - Complete Section III.C. [Committed Effective Dose Equivalent (CEDE)]  Enter the calculated CEDE in the space provided to the left |
| SUM | TEDE (mrem) | - Add the DDE and CEDE values to determine the TEDE (DDE + CEDE = TEDE)  - Enter the TEDE value in the space provided to the left Note: To demonstrate compliance with 64E-5.312(1)(a), TEDE must be < 100 mrem in a year |
|  |  | Method 2. |
|  | DDE (mrem) | - Complete Section III.B. [Deep Dose Equivalent (DDE)]  - Enter the DDE value in the space provided to the left Note: To demonstrate compliance w/ 64E-5.313(2)(b).2, DDE must be < 50 mrem in a year |
|  | Annual average effluent concen-trations < values listed in ALIs, DACs, and Effluent Concentrations* | - Attach documentation for effluent concentrations released in a 12 month period for all radionuclides possessed; the values must be < values listed in ALIs, DACs, and Effluent Concentrations* - Attach the calculations and assumptions for each radionuclide used to determine effluent concentrations; e.g., ventilation rates or volatility based on usage (heating or aerosol production, etc.) |

III.	Annual Public Dose Limit

B.	Deep Dose Equivalent

Tables 1 - 4 describe four DDE calculation methodologies.  Indicate the method used to calculate the DDE value.

- New licensee applicants:	Mark the box indicating the procedure that will be used.

- Current licensees:	Mark the box indicating the procedure that was used.

Following determination of the DDE value, enter it in the appropriate tables provided in Section III.A.  Attach documentation of all assumptions and calculations used to make the DDE dose determination.  Where appropriate, attach an annotated facility diagram to illustrate the radiological conditions present in the workplace.

Method 1.	Occupational Worker Dosimetry Data

If the highest annual dose received by the maximally exposed occupational radiation worker is less than 100 millirem (mrem), it can be assumed that no MOP is likely to receive 100 mrem in a year from the same operations.

Assign a personnel monitoring (PM) device (film badge, TLD, or OSLD) to all occupational radiation workers, or at least to those likely to receive the highest exposures.  The monitoring period should cover at least 12 continuous months of typical RAM use.  In Table 1, record the monitoring period start and end dates, and enter the cumulative dose for the worker receiving the highest dose during the monitoring period.

| Table 1. Occupational Worker Dosimetry Data |  |  |
| --- | --- | --- |
|  | DDE (millirem) | Monitoring Period Start Date: ________________ End Date: _________________ Enter the highest individual cumulative external dose for the 12 month monitoring  period in the space provided to the left. Use this value in Section III.A. |

Method 2.	Dosimetry Data for the Maximally Exposed Individual Member of the Public

If the highest annual dose received by the maximally exposed MOP is less than 100 mrem, it can be assumed that no other MOP is likely to receive 100 mrem in a year from the same operations.

Determine which MOP is likely to receive the highest dose from licensed activities.  Monitor the MOP’s exposures for at least 12 continuous months by assigning a PM device to the individual.  In Table 2, record the monitoring period start & end dates.  Sum the individual’s cumulative dose for the period and enter it in the space provided.  Attach an annotated diagram of the facility to illustrate the radiological conditions present in the workplace.

III.	Annual Public Dose Limit

B.	Deep Dose Equivalent     (Continued)

| Table 2. Dosimetry Data for the Maximally Exposed Individual MOP |  |  |
| --- | --- | --- |
|  | DDE (millirem) | Monitoring Period Start Date: _________________ End Date: _________________ Check to indicate that a facility diagram is attached that identifies all restricted areas and adjacent unrestricted areas, and where the monitored MOP’s workstation is located  Enter the highest cumulative individual dose for the 12 month monitoring period in the space provided to the left. Use this value in Section III.A. |

Method 3.	Environmental Monitoring Data

If environmental monitoring demonstrates that continuous exposure to ambient radiation levels in the workplace for a year results in doses < 100 mrem, then no MOP is likely to exceed the 100 mrem annual public dose limit due to the licensed operations generating the radiation levels.  If environmental monitoring indicates that continuous occupancy would exceed the public dose limit, then occupancy factors may be used to demonstrate compliance.

Post one or more environmental TLDs in the unrestricted areas adjacent to restricted areas, or in the restricted area on a wall adjacent to unrestricted areas, for at least 12 months.  Post badges where the highest radiation exposure is expected and where exposure to non-regulated sources of radiation (e.g., medical patients injected with radionuclides) will not contribute to the measurements.  Record the monitoring period start and end dates in Table 3.  If the results for the 12 month monitoring period total
< 100 mrem, mark Box A, indicating that continuous occupancy was used for the dose determination (24 hours/day, 365.25 days/year = 8,766 hours), and enter the total value in the DDE box provided in Table 3.

If the results for the 12 month monitoring period total > 100 mrem, it may be possible to demonstrate compliance with the annual dose limit by applying a more realistic (but still very conservative) occupancy factor, such as 2000 hours for a work year (8 hour work day, 40 hours a week, 50 weeks a year = 2,000 hours).  Mark Box B if using a normal work week occupancy factor to calculate the DDE.

Example:  The total dose measured by the environmental badges = 280 mrem
The dose received by a MOP working 2,000 hours in the area that the badge was posted is

280 mrem/8,766 hrs  =  .032 mrem/hr  x  2,000 hrs  =  64 mrem

Using a 2000 hour occupancy factor means that any annual dose from environmental monitoring that totals < 438 mrem will demonstrate compliance

Example:  438 mrem/8,766 hrs  =  .049 mrem/hr  x  2,000 hrs  =  99.9 mrem

If the results for the 12 month monitoring period total > 438 mrem, compliance may still be demonstrated by using an even more realistic occupancy factor, provided the number can be legitimized by supporting documentation (e.g., employment records).

Example:  Environmental badges total 680 mrem for the 12 month monitoring period; time sheets indicate that a conservative estimate of the most time spent by any MOP in the monitored area is 25 hours a week, 50 weeks a year = 1,250 hours.

680 mrem/8,766 hrs  =  .078 mrem/hr x 1,250 hrs  =  97 mrem

III.	Annual Public Dose Limit

B.	Deep Dose Equivalent     (Continued)

In each case, attach an annotated diagram of the facility to illustrate the radiological conditions present in the workplace and to indicate the location of posted badges.

Notes:	1.  Protect posted badges from adverse environmental conditions such as excessive 		 heat and light.

2.  Only specifically designed badges are acceptable for environmental monitoring; PM 
		     badges are not appropriate.  Contact the Bureau of Radiation Control for more 
		     information.

| Table 3. Environmental Monitoring Data |  |  |
| --- | --- | --- |
|  | DDE (millirem) | Monitoring Period Start Date: ________________ End Date: __________________ A. Check if calculations are based on continuous year-round occupancy (8766 hours) in unrestricted areas OR B. Check if calculations are adjusted for workplace occupancy factors (e.g., 2000 hours for a work year) in unrestricted areas ******************************************************************************************************************************* Check to indicate that a facility diagram identifying restricted areas, adjacent unrestricted areas, and the location of posted environmental monitor is attached  Enter the environmental monitor’s highest cumulative dose for the 12 month monitoring period. Use this value in Section III.A. |

Method 4.	Radiation Level Data

Survey measurements and calculations can be used to demonstrate that the radiation levels resulting from licensed operations are not likely to cause any MOP to exceed the annual public dose limit.

Radiation levels generated by RAM present in the workplace can be determined by direct measurement with survey instruments, or from indirect information, such as radioactive material package transport index values (describing radiation levels at 1 meter from a package’s exterior surface).  The radiation level data can then be used with the inverse square law to calculate the DDE.  In Table 4, check to indicate use of either radiation survey instrument measurements (Box A-1) or RAM package Transport Index (TI) values (Box B-1) with the inverse square law to calculate the DDE.

The issue of occupancy factors is addressed by selecting one of two options provided in Table 4.  Check off the Box A-2 to indicate use of the most conservative scenario -- assuming a MOP is continuously present in the unrestricted area (24 hours/day, 365.25 days/year = 8766 hours).  Check Box B-2 to indicate use of a more realistic (but still very conservative) assumption -- the individual located in the unrestricted area is present during all business hours (8 hours/day x 40 hours/week x 50 weeks/year = 2,000 hours).

III.	Annual Public Dose Limit

B.	Deep Dose Equivalent     (Continued)

| Table 4. Radiation Level Data |  |  |
| --- | --- | --- |
|  | DDE (millirem) | A-1. Check to indicate use of radiation survey instrument measurements and the inverse square law to calculate the DDE OR B-1. Check to indicate use of RAM package Transport Index (TI) values or RAM package surface radiation levels and the inverse square law to calculate DDE *************************************************************************************************************************************** A-2. Check if dose is based on continuous year-round occupancy (8766 hours) in unrestricted areas OR B-2. Check if dose has been adjusted for workplace occupancy factors (e.g., 2000 hours for a work year) in unrestricted areas *************************************************************************************************************************************** Check to indicate that documentation of all calculations is attached, along with instrument identification, specifications and calibration information Check to indicate a facility diagram showing restricted and unrestricted areas is attached  Enter the calculated DDE in the space provided to the left; use this value in Section III.A. |

Inverse Square Law:    I2  =

| I1 R12 | Where: | I1 = intensity (radiation dose rate) at distance R1 |
| --- | --- | --- |
| R22 |  | I2 = intensity (radiation dose rate) at distance R2. |
|  |  | R1 = distance from RAM with dose rate I1 |
|  |  | R2 = distance from RAM where dose rate I2 is calculated |

Notes:		A.  This formula has 2 limitations: (a) it only applies to gamma-emitting point sources; 				      and (b) the closest distance should be at least 5 source diameters.

B.  If using transport package exterior radiation levels, set R1 = 1 inch.

Example of an Inverse Square Law Calculation Using Survey Meter Measurements

A lab contains a variety of sealed sources.  The sources may be treated as a single point source by positioning them together for the measurement.  Assuming a collective source diameter of 12 inches, a radiation measurement (I1) is taken at a distance equal to at least 5 source diameters (R1) from the grouped sources, shielded behind lead brick corral.  The intensity (I2) at 10 feet (R2) is the unknown value being sought (the distance to the nearest unrestricted area).

I1 = 0.1 mR/hr				         0.1 x (60)2	A 2,000 hour occupancy factor yields:
I2 = ? mR/hr				I2  =  	 (120)2			0.025 mR/hr  x  2,000 hours
R1 = 60 in. (5 x 12 in.)						=  50 mrem 
R2 = 120 in. (10 ft.)			I2  =  .025 mR/hr		=  DDE

III.	Annual Public Dose Limit

B.	Deep Dose Equivalent     (Continued)

Example of an Inverse Square Law Calculation Using a Package Transport Index

A shipping case used to store a portable nuclear density gauge bears a Radioactive Yellow II label that shows its TI = 1.2.  The nearest MOP workstation is located 24 feet away.

I1   = 1.2 mR/hr		         1.2 x (3.3)2		A 2,000 hour occupancy factor yields:
I2   = ? mR/hr			I2  =  	  (24)2			0.023 mR/hr  x  2,000 hours
R1 = 3.3 feet (1 meter)					=  46 mrem
R2 = 24 feet			I2  =  0.023 mR/hr		=  DDE

C.	Committed Effective Dose Equivalent

If licensed for, or seeking licensure for use of unsealed radioactive material (RAM) or both sealed and unsealed RAM, the internal as well as external radiation hazard must be evaluated to demonstrate compliance with the public dose limits described section 64E-5.312, Florida Administrative Code (F.A.C.)

Committed Effective Dose Equivalent (CEDE) refers to the dose resulting from internal radiation exposures.  The CEDE is combined with the Deep Dose Equivalent (DDE), the dose from external whole body exposures, to produce the Total Effective Dose Equivalent (TEDE), the dose resulting from internal and external radiation exposures.  Refer to section 64E-5.101, F.A.C., for complete definitions of these terms.

This procedure provides a method of calculating the CEDE value required by Section III.A.

Note:    The following may be excluded from CEDE calculations:

   Sealed sources

   Exposure from oral pathways or wounds

   Molybdenum 99 from Mo-99/Tc-99m generators based on the 
		  Mo-99 breakthrough limit of < 0.15 Ci of Mo-99 per mCi of 			  Tc-99m

- If any current or requested RAM use does not correspond to the listed options, perform separate calculations of their CEDE dose contributions; attach a description of each RAM type and quantity, and the CEDE calculations performed to determine their dose contribution.

- Sum the applicable mrem values from the marked boxes and enter the sum in the last box; use this value in Section III.A.

III.	Annual Public Dose Limit

C.	Committed Effective Dose Equivalent    (Continued)

| Description of Radioactive Materials Use -- Types & Quantities |  |  |
| --- | --- | --- |
|  | 5 mrem | For any RAM with an ALI value > 100 Ci, total use < 400 mCi in any 12 month period, except C-14; this includes the following RAM: Cl-36 Cu-64 H-3 N-65 P-32 Rb-81m Sc-46 Co-57 Fe-55 Hg-203 Na-22 P-33 S-36 Sm-153 Co-60 Fe-59 I-123 Ni-63 Rb-81 Sb-119 Zn-65 Note: ALI values are listed in Table I, Column 2 of ALIs, DAC, and Effluent Concentrations, July 1993 (appended to the end of Chapter 64E-5. F.A.C.) |
|  | 1 mrem | C-14 use in any form is < 400 mCi in any 12 month period |
|  | 1 mrem | C-14 use in non-volatile forms is < 4 Ci in any 12 month period |
|  | 1 mrem | I-125 use in any form is < 1.2 mCi in any 12 month period |
|  | 1 mrem | I-125 use in non-volatile forms (other than NaI and not involving heating or exothermic chemical reaction) is < 1200 mCi in any 12 month period |
|  | 1 mrem | I-125 use in non-volatile forms (other than NaI and involving heating or exothermic chemical reaction) is < 120 mCi in any 12 month period |
|  | 1 mrem | I-125 use as gases and volatile forms (other than NaI and not involving heating or exothermic chemical reaction) is < 12 mCi in any 12 month period |
|  | 1 mrem | I-125 use as gases and volatile forms (other than NaI, and involving heating or exothermic chemical reaction) is < 1.2 mCi in any 12 month period |
|  | 1 mrem | I-131 use in any form is < 1 mCi in any 12 month period |
|  | 1 mrem | I-131 use in non-volatile forms (other than NaI, and not involving heating or exothermic chemical reaction) is < 1000 mCi in any 12 month period |
|  | 1 mrem | I-131 use in non-volatile forms (other than NaI, and involving a heating or exothermic chemical reaction) is < 100 mCi in any 12 month period |
|  | 1 mrem | I-131 use as gases and volatile forms (other than NaI, and not involving heating or exothermic chemical reaction) is < 10 mCi in any 12 month period |
|  | 1 mrem | I-131 use as gases & volatile forms (other than NaI, and involving heating or exothermic chemical reaction) is < 1 mCi in any 12 month period |
|  | _ mrem | Other RAM not listed above; attach description of types, quantities and the calculations performed to determine their CEDE dose contribution |
| SUM | CEDE (mrem) | Sum the applicable doses and enter the calculated total CEDE value in the  space provided to the left; use this value in Section III.A. |