To address requirements specified in Chapter 64E-5, Florida Administrative Code, and by the U.S. Department of Transportation (DOT) in Title 49, Code of Federal Regulations (49 CFR), the procedures described below will be followed.

Only personnel qualified as hazmat employees in accordance with 49 CFR Part 172 requirements are allowed to perform the transport-related functions described below.

I.	Ordering and Receipt

A.	The Radiation Safety Officer (RSO) will place or approve all orders for radioactive material and ensure that the requested material and form are authorized by the license and will not exceed possession limits specified in the license.

B.	Transportation carriers must be provided instructions describing where to deliver packages containing radioactive materials.

II.	Opening Packages

A.	Visually inspect each package for signs of damage (e.g., wet or crushed).  Using a calibrated survy instrument, monitor the external surfaces for contamination of all packages  labeled as White I, Yellow II or Yellow III.  If any damage or excessive radiation levels are noted, immediately notify the RSO or RSO designee.

B.	Open the package with the following precautionary steps:

- Remove the packing slip.

- Open the outer package following the supplier’s instructions when provided.

- Open the inner package and verify that the contents agree with the packing slip.

- Check the integrity of the final source container.  Look for broken seals or vials, loss of liquid, condensation, or discoloration of the packing material.

- If anything is other than expected, stop and notify the RSO.

C.	Wipe the external surface of the final source container and remove the wipe sample to a low background area and monitor it.  If the instrument indicates a reading above background, stop the procedure and notify the RSO.

D.	Confirm that the material received is the material that was ordered.

E.	Monitor the packing material and the empty packages for contamination, before discarding.

- If contaminated, treat this material as radioactive waste.

- If not contaminated, remove or obliterate the radiation labels before discarding in non-radioactive trash.

F.	Records of package opening survey results are maintained for 3 years as specified in section 64E-5.336, F.A.C.

G.	Section 64E-5.327, F.A.C., allows certain exemptions from package contamination monitoring for radioactive material in the form of gas, special form or for the quantities of radioactive material less than or equal to the A1 or A2 quantities defined in Part XV.

III.	Transportation

A.	Markings and labels on radioactive materials transport containers must be durable, legible, in English, and printed on or affixed to the package surface (e.g., a label, tag or sign).  Required package labels and markings cannot be obscured by any other markings, labels or obstructions.

- Required markings include:

- Shipping name;

- RQ (Reportable Quantity; applies to shipments with activities specified
in 49 CFR 172.101, Appendix 2, Table A);

Identification number; and

## RQ

### RQ
 Package
Marking

- Package type.

2.	Required labels include:

| - “Cargo Aircraft Only” label (for shipments by air) - Two DOT warning labels applied to opposite sides of the package, listing the radionuclide and activity in SI units (English units may be listed after SI units) and the package’s Transport Index (TI), the dimensionless number indicating the package’s radiation level at 1 meter. | ##### Cargo Only Label |
| --- | --- |

Package Labeling Criteria

| Warning Label | Max. Rad. Level at Package Surface (mR/hr) | Max. Rad. Level at 1 m (TI) |
| --- | --- | --- |
| RADIOACTIVE WHITE I | 0.5 | None |
| RADIOACTIVE YELLOW II | 50 | 1 |
| RADIOACTIVE YELLOW III | 200 | 10 |

| RADIOACTIVE WHITE I | RADIOACTIVE YELLOW II | RADIOACTIVE YELLOW III |
| --- | --- | --- |

III.	Transportation

B.	Packaging, Marking and Labeling Exceptions for Limited Quantity Shipments

49 CFR Section 173.421 states that if the quantity of radioactive material does not exceed the amount specified in 49 CFR Section 173.425 and conforms to the requirements specified in 49 CFR Section 173.421 it is exempted from the specifications of packaging, labeling and marking (except for the UN identification number marking requirement described in 49 CFR Section 173.422[a]).  The following procedure applies to qualifying shipments of unused radioactive material and waste.

1.	Ensure that the radioactive waste being returned does not exceed the specified limits for “Limited Quantity Shipments.”

2.	Determine that the radiation level at any point on the surface of the package does not exceed 0.5 millirem/hour by monitoring the package prior to the shipment.

3.	Determine that the non-fixed (removable) radioactive surface contamination on the external surface of the package does not exceed the limits specified in 
49 CFR subsection 173.443(a), for example, 22 dpm/cm2 when wiped over a 300 cm2 area.

4.	If the package does not exceed the limits in 1, 2 and 3, the package may be shipped as a limited quantity shipment.  The outside of the inner package or if there is no inner package, the outside of the package itself bears the marking “Radioactive.”

5.	The following statement will be enclosed with the package, included with the packing list or otherwise forwarded with the package:  “The package conforms to the conditions and limitation specified in 49 CFR Section 173.421 for radioactive material, excepted package-limited quantity of material, UN 2910.”

LIMITED SHIPMENT QUANTITES FOR EACH COMMONLY USED RADIONUCLIDES

| Radioisotope | Limited Shipment Quantity (mCi) A2 x 10-4 | Radioisotope | Limited Shipment Quantity (mCi) A2 x 10-4 |
| --- | --- | --- | --- |
| Ba-133 | 8.1 | I-131 | 1.9 |
| Cr-51 | 81 | P-32 | 1.4 |
| C-14 | 8.1 | P-33 | 2.7 |
| Cs-137 | 1.6 | Ra-226 | 0.0081 |
| H-3 | 110 | S-35 | 8.1 |
| Hg-203 | 2.7 | Sn-113 | 5.4 |
| I-125 | 8.1 | Zn-65 | 5.4 |

When shipping more than one type of radioactive material in the same package, the limit on the radioactivity that may be shipped is determined by the lowest curie quantity assigned for items shipped.  The above values have been calculated using information from 49 CFR section 173.425, Table 7 – Activity for Limited Quantities, Instruments and Articles, and 
49 CFR 173.435, Table of A1 and A2 Values for Radionuclides.  The values identified in these tables are may also used for determining the limited shipment quantities of special form sources and limited quantity instruments or articles.

III.	Transportation

C.	Shipping papers

The information required on a shipping paper (bill of lading) depends on the type of shipment being made, as described below.

1.	Private use shipments are sole use shipments, with all loading, transport and unloading carried out by the licensee’s trained personnel in accordance with the shipper’s instructions, which must accompany the package during shipment.  Private use shipments require a bill of lading with the following information.

- Description of shipment (Include the proper shipping name, RQ [if applicable], identification number, hazard class, package type, name and activity of each radionuclide, category of warning label and Transport Index) (same as for common carrier shipments)

- Emergency response telephone number (24-hour-monitored number of a person knowledgeable about the hazards associated with the source of radiation being shipped)

- Date of shipment

While not a requirement, the name, address and telephone number of the shipper may be included as a security enhancement.

2.	Common carrier shipments (packages offered to third parties for transport) require a bill of lading with the information listed below.  If shipped by air, the carrier will provide a “Dangerous Goods Airbill” form to document required information.  Common carrier shipping papers must have all information typed in.  In addition to the information described above, common carrier shipments require a bill of lading with the following information.

- Name and address of shipper (can be the consignee [the licensee offering the package for shipment] or the consignor [service company shipping the package])

- Description of shipment (RQ [if applicable], proper shipping name, hazard class, identification number, type of package, name and activity of each nuclide, category of labeling and Transport Index)

- Emergency response telephone number (24-hour-monitored number of a person knowledgeable about the hazards associated with the source of radiation being shipped)

- Shipper’s certification (statement certifying that the package has been properly classified, described, packaged, marked and labeled, and is in proper condition for transportation)

- Signature of shipper and date of shipment

#### Additional Statements Required for Air Shipments

- Cargo aircraft statement: “Cargo aircraft only”

- Package dimensions

- Candy-stripe borders

III.	Transportation

C.	Shipping papers    (contd.)

3.	Emergency response information (ERI) provides first responders (i.e., medical, fire and law enforcement personnel) with the information needed to take appropriate action in the event of an emergency.  Drivers are required to have in their possession a separate ERI sheet for each type of radioactive material being transported.

4.	Accessibility, Shipping papers and ERI are immediately accessible to the driver during transport.  That is, the papers are within immediate reach and either readily visible to a person entering the driver's compartment or in a holder mounted to the inside of the door on the driver's side of the vehicle.

D.	Inspection

Prior to shipment, transport containers will be inspected to ensure proper packaging and unimpaired physical condition of the container and its closure devices.  Any defects must be promptly reported to the RSO.  The RSO will tag and remove from use any package found to be defective and ensure their repair or replacement.

E.	Blocking and bracing

Transport containers will be blocked and braced to prevent shifting during normal transportation conditions.  Transport containers will not be transported in a vehicle’s passenger compartment.