I.	Radiation Surveys

A.	At the conclusion of each day of use or receipt, survey all areas with a radiation survey instrument where radioactive materials are received and used.

B.	Survey weekly with a radiation survey instrument all areas where radioactive materials or waste are stored.

C.	Surveys shall be completed with an operable and calibrated instrument adhering to the requirements specified in section 64E-5.1319, F.A.C.

II.	Removable Contamination Surveys

A.	Survey weekly for removable contamination in all areas where radioactive materials or wastes are routinely used and stored.

B.	If the radioactive materials authorized by the license are not detectable with the instruments described above in section I.C. above, a removable contamination survey shall be completed at the end of each day of use of all areas where radioactive materials are routinely used.

C.	Radiation detection instrumentation used for removable contamination measurements will adhere to the requirements specified in section 64E-5.1318, F.A.C.  A detailed description of the instruments and procedures used for measuring contamination,
in accordance with subsection 64E-5.1318(2), F.A.C., is attached.

III.	Personnel Surveys

A.	Prior to leaving the restricted area, personnel shall be monitored for contamination in a low background area.

B.	Surveys shall be completed with a calibrated instrument as described in section 
64E-5.314, F.A.C., and capable of measuring dose rates as low as 0.1 millirem per hour, if applicable.

IV.	Action Levels and Decontamination

A.	The action levels for removable contamination as established in paragraph 
64E-5.1319(2)(c), F.A.C., are as follows:

- 100 dpm per 100 square centimeters of any alpha-emitting radioactive materials not listed in this section;

- 1,000 dpm per 100 square centimeters of any beta- or gamma- emitting radioactive materials not listed in this section;

- 50 dpm per 100 square centimeters of any transuranic;

- 2,000 dpm per 100 square centimeters of uranium; or

- 2,000 dpm per 100 square centimeters of any radioactive material with a half-life of less than 80 hours.

IV.	Action Levels and Decontamination

B.	The action levels for dose rate surveys is 2 millirem/hour.  Measured dose rates exceeding this value will be cause for implementing decontamination procedures.

C.	If removable contamination or dose rate action levels are exceeded, decontamination efforts will be taken as described in emergency procedures.  Following completion of decontamination efforts, the contaminated area will be resurveyed to assess effectiveness of decontamination and to ensure that it is now below the established action levels.

V.	Notification and Records

A.	Notify the Radiation Safety Officer (RSO) if the contamination detected during the surveys exceeds the action levels established in subsection 64E-5.1319(2), F.A.C.

B.	Make a record of required surveys and retain a copy for 3 years.  Records shall include the following information:

- The date of the survey;

- An annotated diagram of each area surveyed;

- Background levels;

- Measured dose rates or removable contamination results will be keyed to the diagram.  Measured dose rates will be expressed in millirem per hour.  Removable contamination results will be expressed in dpm per 100 square centimeters or counts per minute if performed with a radiation survey instrument.  Corresponding action levels established for both measured dosed rates and removable contamination results will be identified on the diagram.

- The serial number and model number of the instrument used to make the survey or analyze the samples; and

- The initials of the person who performed the survey.