Sealed sources used for calibration and reference or contained in devices must be tested at regular intervals to ensure that the radioactive material is secure within its capsule and not leaking.  
Leak test (LT) requirements are specified in section 64E-5.1303, Florida Administrative Code (F.A.C.).

I.	Instrumentation and Sensitivity

Select an instrument that is sufficiently sensitive to detect 0.005 microcurie.  For beta sources, use a proportional flow counter or liquid scintillation counter.  For gamma sources, a NaI crystal with a scaler will be appropriate.  Identify instrumentation in facility description.

II.	Determination of Sample Counting Times

Determine the minimum sample counting times needed to distinguish 0.005 microcurie from the background for each instrument.

Measure the background count rate (Rb) in counts per minute (cpm) and record.

Measure the correction factor using a known National Institute of Standards and Technology (N.I.S.T.) source and record.  Assay a certified check source that has the same isotope as the sealed source being tested.  If a certified check source is not available, it will be necessary to use one with a different isotope that has a similar energy spectrum.

CF = Rst-Rb    Rst = count rate of standard (cpm)
	A(Ci)

Example:  Background is 30 cpm and a 10 Ci source measures 40,030 cpm on the instrument.

CF = 40,030-30 = 4000 cpm/Ci.
	10 Ci

Calculate minimum sample counting time (tms) in minutes for the instrument.

Lower Limit of Detection (LLD)=  4.66      Rb

CF       tms

tms =     4.66      2 Rb

CF(.005)

tms = 868,624 x Rb    (minutes)

CF x CF

III.	Analysis, Reporting and Records

A.	Count each wipe at least tms.

Determine count rate for each sample Rs = Ns (cpm)
	ts

Ns = number of counts		ts = sample counting time

Determine activity as follows:

A(Ci) = Rs-Rb  Record in units of microcuries
	CF

- Continue the same analysis procedure for all wipe samples.

- If the wipe sample activity is 0.005 microcurie or greater, notify the RSO.  The source must be withdrawn from use to be repaired or properly disposed.  A report shall be filed within 5 days of the test with the department.

- Sign and date the list of sources, data, and calculations.

IV.	Calculations

Background is 150 counts in 5 minutes or 150 = 30 cpm
	5

10 Ci cesium standard measures 40,030 cpm

CF = 40,030-30 = 4000 cpm/Ci
	10 Ci

CF = 4000 cpm/Ci     Rb = 30 cpm

tms =   868,624 x 30 = 1.63 minutes
	(4000)(4000)

Must count at least 1.63 minutes.

Have chosen to count each sample 5 minutes.

Wipe #1 159 counts in 5 minutes		Wipe #2 164 counts in 5 minutes.

R1 =  159 = 31.8 cpm				R2 =  164 = 32.8 cpm
	5						5

A1 = 31.8-30 = 0.00045 Ci			A2 =  32.8-30 = 0.0007 Ci
	4000						4000

Both are < 0.005 microcurie.