Sections 64E-5.103, .215, .328, .330, .331, and .340, Florida Administrative Code (F.A.C.), address the transfer and disposal of radioactive material.  Licensed radioactive material will be transferred or disposed in accordance with these regulatory requirements.

I.	Disposal Methods

A.	Effluent and Evaporative Release

Liquids may be disposed by release to the sanitary sewer or evaporative release 
in accordance with the regulatory requirements.

- Radioactive material discharged by release into the sanitary sewerage system will be readily soluble in water or is readily dispersible biological material.  Radioactive material released into the sewer in 1 month divided by the average monthly volume of water released into the sewer will not exceed the concentration listed in Table III of the ALIs, DACs, and Effluent Concentrations.  Applicable calculations will be maintained demonstrating compliance with the requirements specified in section 64E-5.330, F.A.C.

Records will be maintained indicating the date, radionuclide, estimated activity 
of the release (in millicuries or microcuries), and the sink where the material is released.

- Dose limits from effluents to unrestricted areas to members of the public will be maintained as required by 64E-5.312, F.A.C., and Table II of the ALIs, DACs, and Effluent Concentrations.  These limits apply at the boundary of the restricted area.  Records will be maintained indicating the date, radioisotope, estimated activity released (in millicuries or microcuries), estimated concentration, and the vent site at which the material is released.

- Liquid scintillation-counting media containing 0.05 microcurie per gram of hydrogen-3 (H-3) or carbon-14 (C-14) may be disposed of without regard to its radioactivity.  Records will be maintained indicating the date, radionuclide, estimated activity (in millicuries or microcuries), calculated concentration in microcuries per gram, and how the material is disposed.

B.	Disposal by Decay-in-Storage (DIS)

Radioactive material, other than special form (sealed) sources, having a half-life less than 90 days may be disposed by DIS.

- Radioactive material with a physical half-life less than 90 days will be segregated according to half-life, when disposed by DIS.

- One container for all waste for DIS - or -separate containers for different types of waste for DIS (e.g., one container for spent in-vitro kits or planchets, a second container for contaminated gloves and absorbent paper, etc., and a third container for expired reference solutions) will be used.

- Waste will be surveyed with all shielding removed, including any shielding provided by the container.

I.	Disposal Methods

B.	Disposal by Decay-in-Storage (DIS)    (Continued)

- When a waste container is full, it will be sealed with string or tape and an identification tag will be attached.  The tag will include the date sealed, the longest-lived radioisotope in the container, and the initials of the person sealing the container.  The container will by transferred to a DIS area.

- Radioactive waste will be held for decay for at least 10 half-lives.

- Radioactive waste will be monitored prior to disposal as in-house trash, according to the following procedure.

a.	Check the radiation detection survey meter for proper operation.

b.	Remove any shielding from around the waste container.

c.	Monitor the waste container in a low background radiation area 
(less than 0.05 millirem per hour).

d.	Monitor all surfaces of each individual waste container.

e.	Monitor waste such that the container does not provide any radiation shielding.

f.	Radioactive waste may be discarded as in-house waste only if it cannot be distinguished from background.  Make a record of the date that the waste container was sealed, the radioactive waste disposal date, and the type of material (e.g., paraphernalia, unused reference or calibration sources).  At the time of disposal, remove or obliterate all radiation labels on the waste container.

g.	Radioactive waste that can be distinguished from background radiation levels is returned to the storage area for further decay or transferred for burial at a low-level waste disposal facility.

- A record of all decay-in-storage radioactive material will be retained for 3 years.  This record will include the date of the disposal; the date on which the radioactive material was placed in storage; the radioisotopes disposed; the model and serial number of the radiation survey instrument used; the background dose rate; the radiation dose rate measured at the surface of each waste container; and the name of the individual who performed the disposal.

C.	Low Level Radioactive Waste Transfer for Land Disposal

Except for material suitable for DIS and some animal carcasses, solids will be transferred to a licensed land disposal facility.  Follow the packaging instructions received from the transfer agent and the land disposal facility operator.  The consignment sheet supplied by the transfer agent will be maintained as the record of disposal. The procedures described in section 64E-5.332, F.A.C., will be followed.

I.	Disposal Methods

D.	Generally Licensed In-Vitro Kits Released to In-House Waste

Waste from in-vitro kits that are generally licensed are exempt from waste disposal regulations.  Radioactive labels will be defaced or removed.  No record of release or radiation measurement will be maintained.

II.	Return to Manufacturer or Transfer to a Licensed Recipient

Licensed radioactive material may be returned to the manufacturer for disposal or transferred to a licensed recipient.  Prior to transferring or disposing of a radioactive source, obtain documentation of the transferee’s authorization to possess the radioactive material.  Retain either a copy of the other entity’s radioactive materials license, or one of the other verification methods listed in subsection 64E-5.215(4), F.A.C., as evidence of an authorized transfer.

As a minimum, documentation of the transfer includes the following:

- The material being transferred (source manufacturer name, model and serial number, type and activity of radioactive material, as applicable);

- The date of the transfer;

- The name, address, and license number of the transferor and transferee; and

- The signatures of the individuals shipping and/or receiving the radiation source.

Form DH-1059 “Certificate – Disposition of Radioactive Materials” may be used to document the transfer/disposal of radioactive material; the form is available on the Bureau of Radiation Control website at http://www.doh.state.fl.us/environment/radiation.  All transfer and disposal records are retained on file until license termination.